## bmHPCrawData
atlasPartsId | 20220170200173
manufacturerSerialNo | SP0553
eventDescription | PC_HYBRID
date [dd/mm/yyyy] | 20/08/2002
location [instituteCode(DB)] | 017
personInitial | YU
problem [YES/NO] | NO
pass [YES/NO] | YES
comments
C1~C28 capacitorsLotID | IA1D10J39
C51~C58 capacitorsLotID | IA1N29M26
C71~C75 HV capacitorsLotID | IA1D13Y04
R27~R30 resitorsLotID | C9Y0401
R33~R34 resitorsLotID | C9X2913
R35 resitorsLotID | N9X0857
TM1~TM2 thermistersLotID | 0013
CON connectorLotID | 2001
FLEX_CIRCUIT LotID | 016-006
PA pitch adaptorsLotID | LOT20020911Al1um1/2
CC carbon-carbon bridgesLotID | 3rd20020124
Conductive epoxyLotID | 0221011
Thermal epoxyLotID | 1225011
solderBallLotID | SS029966,SS021036
solderReelLotID | 01.12.27
FLEX_CIRCUIT visual inspection[G/NG] | G
Integrity test of lines [G/NG] | G
dateCCbridge attaching [dd/mm/yyyy] | 20/08/2002
visual inspection [G/NG] | G
Thickness Link0 (<0.65) | 0.615
Thickness Link1 (<0.65) | 0.609
Bow (long) Link0 (< 0.075) | -0.087
Bow (long) Link1 (< 0.075) | -0.043
Bow (across) Link0 (< 0.075) | -0.071
Bow (across) Link1 (< 0.075) | -0.041
Twist Link0 (<0.1) | 0.065
Twist Link1 (<0.1) | 0.003
dateComponent soldering [dd/mm/yyyy] | 22/08/2002
visual inspection [G/NG] | G
number of Wirebond-Pull on PC hybrid | 4
average (>6.0gr) | 8.12
range(max-min) [gr] | 0.73
min [gr] | 7.64
datePitchAdaptor attaching [dd/mm/yyyy] | 18/09/2002
visual inspection [G/NG] | G
number of Wirebond-Pull on PA (Lot) | 10
average (>6.0gr) | 10.72
range(max-min) [gr] | 2.05
min [gr] | 9.54
Capacitance Vcc-GND (microF) at 1KHz (4.0<X<4.8) | 11.54
Capacitance Vdd-GND (microF) at 1KHz (4.0<X<4.8) | 11.72
Impedance HV-GND (kOhm) at 100Hz ( 30<X< 36) | 30.60
Impedance HV-GND (kOhm) at 1KHz ( 9<X< 11) | 9.91
Impedance HV-GND (kOhm) at 10KHz (5.7<X<7.7) | 6.63
Impedance HV-GND (kOhm) at 100KHz (5.0<X<7.1) | 6.21
Resistance R27 (Ohm) ( 99<X<101) | 100.74
Resistance R28 (Ohm) ( 99<X<101) | 100.78
Resistance R29 (Ohm) ( 99<X<101) | 100.61
Resistance R30 (Ohm) ( 99<X<101) | 100.62
Resistance TM1 (kOhm) at 25 deg.C. ( 9.8<X<10.2) | 10.02
Resistance TM2 (kOhm) at 25 deg.C. ( 9.8<X<10.2) | 9.83
LV leakage current for Icc (nA) at 10V (X<10) | 48.71
LV leakage current for Idd (nA) at 10V (X<10) | 47.87
HV leakge current (nA) at 500V (X<10) | 8.41
dateTC [dd/mm/yyyy] | 29/08/2002
TC loTemperature [C]
TC hiTemperature [C]
TC repetition
visual inspection [G/NG] | G
Thickness Link0 (<0.65) | 0.614
Thickness Link1 (<0.65) | 0.605
Bow (long) Link0 (< 0.075) | -0.065
Bow (long) Link1 (< 0.075) | -0.037
Bow (across) Link0 (< 0.075) | -0.074
Bow (across) Link1 (< 0.075) | -0.046
Twist Link0 (<0.1) | 0.071
Twist Link1 (<0.1) | -0.016
Capacitance Vcc-GND (microF) at 1KHz (4.0<X<4.8) | 11.43
Capacitance Vdd-GND (microF) at 1KHz (4.0<X<4.8) | 11.59
Impedance HV-GND (kOhm) at 100Hz ( 30<X< 36) | 30.80
Impedance HV-GND (kOhm) at 1KHz ( 9<X< 11) | 9.91
Impedance HV-GND (kOhm) at 10KHz (5.7<X<7.7) | 6.64
Impedance HV-GND (kOhm) at 100KHz (5.0<X<7.1) | 6.21
Resistance R27 (Ohm) ( 99<X<101) | 100.74
Resistance R28 (Ohm) ( 99<X<101) | 100.78
Resistance R29 (Ohm) ( 99<X<101) | 100.61
Resistance R30 (Ohm) ( 99<X<101) | 100.62
Resistance TM1 (kOhm) at 25 deg.C. ( 9.8<X<10.2) | 10.03
Resistance TM2 (kOhm) at 25 deg.C. ( 9.8<X<10.2) | 9.84
LV leakage current for Icc (nA) at 10V (X<10) | 48.21
LV leakage current for Idd (nA) at 10V (X<10) | 46.63
HV leakge current (nA) at 500V (X<10) | 8.13
visual inspection [G/NG] | G
Height of Component C73 | 2.236
Height of Component C53 | 2.356
Height of Component C54 | 2.36
Height of Component C74 | 2.236
Height of Component C55 | 2.364
Height of Component C56 | 2.349