## bmHPCrawData
atlasPartsId | 20220330200054
manufacturerSerialNo | SP0660
eventDescription | PC_HYBRID
date [dd/mm/yyyy] | 28/09/2002
location [instituteCode(DB)] | 017
personInitial | YU
problem [YES/NO] | NO
pass [YES/NO] | YES
comments
C1~C28 capacitorsLotID | IA1D10J39
C51~C58 capacitorsLotID | IA1N29M26
C71~C75 HV capacitorsLotID | IA1D13Y04
R27~R30 resitorsLotID | C9Y0401
R33~R34 resitorsLotID | C9X2913
R35 resitorsLotID | N9X0857
TM1~TM2 thermistersLotID | 0013
CON connectorLotID | 2001
FLEX_CIRCUIT LotID | 003-001
PA pitch adaptorsLotID | 26
CC carbon-carbon bridgesLotID | 2nd.2001.12.27
Conductive epoxyLotID | 0221011
Thermal epoxyLotID | 1225011
solderBallLotID | SS029966,SS021036
solderReelLotID | 110404642W31C
FLEX_CIRCUIT visual inspection[G/NG] | G
Integrity test of lines [G/NG] | G
dateCCbridge attaching [dd/mm/yyyy] | 28/09/2002
visual inspection [G/NG] | G
Thickness Link0 (<0.65) | 0.613
Thickness Link1 (<0.65) | 0.616
Bow (long) Link0 (< 0.075) | 0.027
Bow (long) Link1 (< 0.075) | 0.013
Bow (across) Link0 (< 0.075) | -0.047
Bow (across) Link1 (< 0.075) | -0.029
Twist Link0 (<0.1) | 0.034
Twist Link1 (<0.1) | 0.007
dateComponent soldering [dd/mm/yyyy] | 1/10/2002
visual inspection [G/NG] | G
number of Wirebond-Pull on PC hybrid | 4
average (>6.0gr) | 9.79
range(max-min) [gr] | 1.09
min [gr] | 9.33
datePitchAdaptor attaching [dd/mm/yyyy] | 05/11/2002
visual inspection [G/NG] | G
number of Wirebond-Pull on PA (Lot) | 10
average (>6.0gr) | 9.06
range(max-min) [gr] | 1.45
min [gr] | 8.29
Capacitance Vcc-GND (microF) at 1KHz (4.0<X<4.8) | 11.52
Capacitance Vdd-GND (microF) at 1KHz (4.0<X<4.8) | 11.62
Impedance HV-GND (kOhm) at 100Hz ( 30<X< 36) | 30.20
Impedance HV-GND (kOhm) at 1KHz ( 9<X< 11) | 9.90
Impedance HV-GND (kOhm) at 10KHz (5.7<X<7.7) | 6.63
Impedance HV-GND (kOhm) at 100KHz (5.0<X<7.1) | 6.20
Resistance R27 (Ohm) ( 99<X<101) | 100.88
Resistance R28 (Ohm) ( 99<X<101) | 100.81
Resistance R29 (Ohm) ( 99<X<101) | 100.69
Resistance R30 (Ohm) ( 99<X<101) | 100.80
Resistance TM1 (kOhm) at 25 deg.C. ( 9.8<X<10.2) | 10.10
Resistance TM2 (kOhm) at 25 deg.C. ( 9.8<X<10.2) | 10.10
LV leakage current for Icc (nA) at 10V (X<10) | 47.26
LV leakage current for Idd (nA) at 10V (X<10) | 48.64
HV leakge current (nA) at 500V (X<10) | 8.87
dateTC [dd/mm/yyyy] | 15/10/2002
TC loTemperature [C]
TC hiTemperature [C]
TC repetition
visual inspection [G/NG] | G
Thickness Link0 (<0.65) | 0.614
Thickness Link1 (<0.65) | 0.616
Bow (long) Link0 (< 0.075) | -0.035
Bow (long) Link1 (< 0.075) | -0.025
Bow (across) Link0 (< 0.075) | -0.044
Bow (across) Link1 (< 0.075) | -0.030
Twist Link0 (<0.1) | 0.026
Twist Link1 (<0.1) | -0.005
Capacitance Vcc-GND (microF) at 1KHz (4.0<X<4.8) | 11.43
Capacitance Vdd-GND (microF) at 1KHz (4.0<X<4.8) | 11.53
Impedance HV-GND (kOhm) at 100Hz ( 30<X< 36) | 30.40
Impedance HV-GND (kOhm) at 1KHz ( 9<X< 11) | 9.91
Impedance HV-GND (kOhm) at 10KHz (5.7<X<7.7) | 6.64
Impedance HV-GND (kOhm) at 100KHz (5.0<X<7.1) | 6.20
Resistance R27 (Ohm) ( 99<X<101) | 100.88
Resistance R28 (Ohm) ( 99<X<101) | 100.81
Resistance R29 (Ohm) ( 99<X<101) | 100.69
Resistance R30 (Ohm) ( 99<X<101) | 100.80
Resistance TM1 (kOhm) at 25 deg.C. ( 9.8<X<10.2) | 10.11
Resistance TM2 (kOhm) at 25 deg.C. ( 9.8<X<10.2) | 10.11
LV leakage current for Icc (nA) at 10V (X<10) | 48.40
LV leakage current for Idd (nA) at 10V (X<10) | 47.86
HV leakge current (nA) at 500V (X<10) | 8.19
visual inspection [G/NG] | G
Height of Component C73 | 2.263
Height of Component C53 | 2.345
Height of Component C54 | 2.358
Height of Component C74 | 2.220
Height of Component C55 | 2.360
Height of Component C56 | 2.340