## Introduction
I. DATASHEETS (V3.2)
Module datasheets are made of two excel books: main PC_HYBRID-ASIC_HYBRID-MODULE chain, and
a flow of BASEBOARD-SENSOR_BASEBOARD till integrated into the main chain.
main - moduleDataSheets3.2.xls
baseboard - moduleDataSheets3.2SB.xls
II. Datasheet in steps
(SPELL CRITICAL)
|  |  |  | Event description | atlasPartsId | Mandatory(*) | Reference document
1. Components | [5 empty] | (at least initially)
| 1.1 Baseboards |  |  | BASEBOARD | <label on baseboard> | * | SCT-BM-FDR-5.2
| 1.2 Passive-component-loaded Hybrids |  |  | PC_HYBRID | <label on hybrid> | * | SCT-BM-FDR-5.3
2. Assembly
| 2.1 Sensor-baseboard assembly |  |  | SENSOR_BASEBOARD | <label on baseboard> | * | SCT-BM-FDR-7
| 2.2 ASIC-hybrid assembly |  |  | ASIC_HYBRID | <label on hybrid> | * | SCT-BM-FDR-7
| 2.3 Module assembly |  |  | MODULE | <label on hybrid> | * | SCT-BM-FDR-7
3. Tests
3.1 Sensor-baseboard unit
|  | 3.1.1 I-V measurement |  | I_V-SB | <label on baseboard> | * | SCT-BM-FDR-7
|  | 3.1.2 I stability |  | I_STABILITY-SB | <label on baseboard> |  | SCT-BM-FDR-7
|  | 3.1.3 Metrology |  | SURVEY_XY-SB | <label on baseboard> | * | SCT-BM-FDR-7
|  |  |  | SURVEY_Z-SB | <label on baseboard> |  | SCT-BM-FDR-7
3.2 ASIC-hybrid unit
3.2.1 Initial assembly
|  |  | 3.2.1.1 Electrical performance | ELECTRICAL-INITIAL-ASIC_HYBRID | <label on hybrid> | * | Electrical Tests of SCT Hybrids and Modules, Version 3.01,
3.2.2 Long-term test
|  |  | 3.2.2.1 Measurement during test | LONGTERM-ASIC_HYBRID | <label on hybrid> | * | SCT-BM-FDR-7
|  |  | 3.2.2.2 Electrical performance | ELECTRICAL-LT-ASIC_HYBRID | <label on hybrid> | * | Electrical Tests of SCT Hybrids and Modules, Version 3.01,
3.3 Module unit
3.3.1 Initial assembly
|  |  | 3.3.1.1 Metrology | SURVEY_XY-INITIAL | <label on hybrid> | * | SCT-BM-FDR-7
|  |  |  | SURVEY_Z-INITIAL | <label on hybrid> | * | SCT-BM-FDR-7
|  |  | 3.3.1.2 I_V measurement | I_V-INITIAL | <label on hybrid> | * | SCT-BM-FDR-7
|  |  | 3.3.1.3 Electrical performance | ELECTRICAL-INITIAL | <label on hybrid> | * | Electrical Tests of SCT Hybrids and Modules, Version 3.01,
3.3.2 Thermal cycling
|  |  | 3.3.2.1 Metrology | SURVEY_XY-TC | <label on hybrid> |  | SCT-BM-FDR-7
|  |  |  | SURVEY_Z-TC | <label on hybrid> |  | SCT-BM-FDR-7
|  |  | 3.3.2.2 I_V measurement | I_V-TC | <label on hybrid> |  | SCT-BM-FDR-7
|  |  | 3.3.2.3 Electrical performance | ELECTRICAL-TC | <label on hybrid> |  | Electrical Tests of SCT Hybrids and Modules, Version 3.01,
3.3.3 Long-term test
|  |  | 3.3.3.1 Measurement during test | LONGTERM-MODULE | <label on hybrid> | * | SCT-BM-FDR-7
|  |  | 3.3.3.2 I stability durring test | I_STABILITY-LT | <label on hybrid> | * | SCT-BM-FDR-7
|  |  | 3.3.3.3 Metrology | SURVEY_XY-LT | <label on hybrid> | * | SCT-BM-FDR-7
|  |  |  | SURVEY_Z-LT | <label on hybrid> | * | SCT-BM-FDR-7
|  |  | 3.3.3.3 I_V measurement | I_V-LT | <label on hybrid> | * | SCT-BM-FDR-7
|  |  | 3.3.3.5 Electrical performance | ELECTRICAL-LT | <label on hybrid> | * | Electrical Tests of SCT Hybrids and Modules, Version 3.01,
|  | 3.3.4 Irradiation |  |  |  |  | (This is just afer the irradiation)
|  |  | 3.3.4.1 I_V measurement | I_V-IRRAD | <label on hybrid> |  | SCT-BM-FDR-7
|  |  | 3.3.4.2 I_STABILITY | I_STABILITY_IRRAD | <label on hybrid>
|  |  | 3.3.4.3 Electrical performance | ELECTRICAL-IRRAD | <label on hybrid> |  | Electrical Tests of SCT Hybrids and Modules, Version 3.01,
|  | 3.3.5 Irradiation annealled |  |  |  |  | (This is afer the annealing)
|  |  | 3.3.3.2 I_V measurement | I_V-ANNEAL | <label on hybrid> |  | SCT-BM-FDR-7
|  |  | 3.3.3.3 I stability | I_STABILITY-ANNEAL | <label on hybrid> |  | SCT-BM-FDR-7
|  |  | 3.3.5.1 Electrical performance | ELECTRICAL-ANNEAL | <label on hybrid> |  | Electrical Tests of SCT Hybrids and Modules, Version 3.01,
3.3.6 Laser scan
|  |  | 3.3.6.1 Electrical performance | LASER | <label on hybrid> |  | TBD
3.3.7 Source scan
|  |  | 3.3.7.1 Electrical performance | SOURCE | <label on hybrid> |  | TBD
3.3.8 Testbeam
|  |  | 3.3.8.1 Electrical performance | ELECTRICAL-TB | <label on hybrid> |  | Electrical Tests of SCT Hybrids and Modules, Version 3.01,
3.3.9 System test
|  |  | 3.3.9.1 Electrical performance | ELECTRICAL-ST | <label on hybrid> |  | Electrical Tests of SCT Hybrids and Modules, Version 3.01,
4. Module summary |  |  |  | MODULE-SUMMARY | <label on hybrid>
III. OTHERS
1. "location"
Use SCT-Database (DB) definition, which is listed in the instituteCode sheet
ATLAS has defined the "institute code(xxx)" in numeric and mnemonic.
SCTDB also defines both numeric and alphabetic abbreviation where the numeric is the same as the ATLAS while
the alphabetic happens to be different.
SCTDB has already been set and the code has to follow the DB for the database.
2. ELECTRICAL tests
The content of the ELECTRICAL tests is the output generated automatically by the SCTDAQ program.
The sheet of ELECTRICAL shows concisely what is in the output.
3. Database
3.1 Database implementation
Since the visible barcode of the module is the barcode on the pigtail of the hybrid which is attached in the PC_HYBRID,
this barcode will be the "atlasPartsId" of the module.
All the datasheets, except the baseboard, in this book are part of the "atlasPartsId" item, and
implemented in the database under this umbrella.
BASEBOARD and SENSOR_BASEBOARD will carry the "atlasPartsId" of the BASEBOARD and
implemented under this umbrella.
3.2 Uploading sheets
We prepare/fill the datasheets given in this excel file.
The sheets are meant to be simple and straight-forward, to be easy to fill.
One may regard these as REPORT, which indeed helps us in two aspects:
(1) we have already had the REPORT from the database
(2) the sheets help us to analyze the data by our own point of view
The database uploading programme may require a more structured and database-protocol-filled datasheet.
Such formatting is straight-forward from the datasheets we prepare and fill here, and
automatic once the (formatting) excel file/sheet is ready.
Which datasheets to choose?
A proposal is
(1) fill and keep the datasheets of this excel file(as the first object)
(2) a formatting excel file/sheet is to be made, in collaboration with the database manager, (and append to this file?)
so that a grand database-protocoled datasheet is automatically generated
(3) individual sheet is also made possible to be uploaded, as needed
4. Sensor leakage current
The RAWDATA of the sensor leakage current are stored as measured, together with the temperature.
Sensor leakage current has a large temperature dependence.
Comparison is to be made at the temperature of 20 deg.C by using the temperature dependence equation of
I(20[C])=I(T[C])*(293/(273+T[C])^1.5*exp(-7019*(1/293-1/(273+T[C]))
where E_g=1.21 eV is used to obtain the coefficient, 7019.
IV. Version history
| 1.0 |  |  | primitive version
| 2.0 |  |  | second iteration
| 3.0 |  |  | established version
3.1
| 3.2 |  | 07/12/2001 | replace surveyZ sheets with surveyZ...xls version 2.2
prepare all default sheets, such as -INITIAL, -TC, -LT
<atlasPartsId> inherited from PC_HYBRID for all sheets except BASEBOARD
delete R36 line in PC_HYBRID
ASIC id=MFR_SER_NO=siteID-lotID-waferID-Xnn-Ynn in ASIC_HYBRID
use SCT-DB institute code for location

## instituteCodes
ATLAS institute code (Mnemonic and Numeric) and SCT-Database(DB) code
Insitute | Mne | Num | DB
Barrel module:
Dept. of Physics Faculty of Science Okayama Univ. | OKA | 173 | OKA
Hiroshima University | HIR | 016 | HIROSH
High Energy Accelerator Research Organization (KEK) | KEK | 017 | KEK
Kyoto University of Education | EKY | 018 | EKY
University of Tsukuba<br>Institute of Physics | TSU | 174 | TSUKUBA
University of Bergen | BER | 003 | BERGEN
University of Oslo | OSL | 029 | OSLO
University of Uppsala | ISV | 038 | UPPS
Rutherford Appleton Laboratory | RAL | 033 | RAL
University of Birmingham | BIR | 006 | BHAM
University of Cambridge | CAM | 007 | CAMB
University of London<br>Queen Mary&Westfield Coll. | QMW | 022 | QMW
University of London<br>University College London | UCL | 023 | UCL
University of Oxford | OXF | 030 | OXFORD
Lawrence Berkeley National Laboratory (LBNL) | LBL | 004 | LBL
University of California, Santa Cruz | CSC | 037 | UCSC
Forward module:
University of Melbourne | MEL | 025 | MELBRN
Acad Scie Czech Republic (ASCR)<br>Institute of Physics | EPS | 043 | PRAGAS2
Charles University | CHP | 044 | PRAGCU
Czech Technical University of Prague CTU (CVUT) | UTP | 042 | PRAGCVUT
Albert-Ludwigs-Universitat Freiburg | ALU | 012 | FREIBURG
Max-Planck-Institut fur Physik | MPI | 027 | MUNICH
NIKHEF (NIKHEF) | FIA | 028 | NIKHEF
Henryk Niewodniczanski Institute of Nuclear Physics | INP | 010 | CRACINP
Univ Mining and Metallurgy<br>Physics and Nuclear Tech | PNT | 009 | CRACFPNT
Inst. for High Energy Physics, Protovino (IHEP) | HEP | 032 | PROT
Jozef Stefan Institute | LJU | 021 | JSI
Univ Valencia<br>Instituto de Fisica Corpuscular (IFIC) | FIC | 039 | VAL
European Organization for Nuclear Research (CERN) | CER | 008 | CERN
Universite de Geneve | UGE | 013 | GVA
Lancaster University | LAN | 019 | LANCS
The University of Manchester<br>Department of Physics | MAN | 024 | MAN
University of Glasgow | GLA | 014 | GLASGOW
University of Liverpool | LIV | 020 | LPOOL
University of Sheffield | SHE | 034 | SHEFF

## PC_HYBRID
atlasPartsId | 20220170200005
manufacturerSerialNo | K5003
eventDescription | PC_HYBRID
date [dd/mm/yyyy] | 35772
location [instituteCode(DB)] | KEK
personInitial | Y.Ikegami
problem [YES/NO] | YES
pass [YES/NO] | YES
comments | No initail mech. Data
C1~C28 capacitors/batch | 20220170111002
C51~C58 capacitors/batch | 20220170112002
C71~C75 HV capacitors/batch | 20220170113001
R27~R30 resitors/batch | 20220170114001
R33~R34 resitors/batch | 20220170115001
R35 resitors/batch | 20220170116001
TM1~TM2 thermisters/batch | 20220170118001
CON connector/batch | 20220170119001
PA pitch adaptors/batch | 20220170121001
CC carbon-carbon bridges/batch | 20220170122001
Conductive epoxy/batch | 20220170131001
Thermal epoxy/batch | 20220170132001
Attaching CCbridges to flex circuit/batch | 20220170133001
Stuffing passive components/batch | 20220170141001
FLEX_CIRCUIT visual inspection | good
Integrity test of lines | good
BRIDGE visual inspection | good
Thickness Link0 (<650)
Thickness Link1 (<650)
Bow (long) Link0 (< 75)
Bow (long) Link1 (< 75)
Bow (across) Link0 (< 75)
Bow (across) Link1 (< 75)
Twist Link0 (<100)
Twist Link1 (<100)
PASSIVE_COMPO visual inspection | good
number of Wirebond-Pull | 4
Wire-bond pull strength (>6.0gr) | 9
Wire-bond pull max-min [gr] | 1.7
Wire-bond pull min [gr] | 8.3
PITCH_ADAPTOR visual inspection | good
Capacitance Vcc-GND (microF) at 1KHz (4.0<X<4.8) | 11.54
Capacitance Vdd-GND (microF) at 1KHz (4.0<X<4.8) | 11.51
Impedance HV-GND (kOhm) at 100Hz ( 30<X< 36) | 33
Impedance HV-GND (kOhm) at 1KHz ( 9<X< 11) | 10
Impedance HV-GND (kOhm) at 10KHz (5.7<X<7.7) | 6.7
Impedance HV-GND (kOhm) at 100KHz (5.0<X<7.1) | 6.08
TC loTemperature [C] | -30 | -30
TC hiTemperature [C] | 50 | +50
TC repetition | 5 | 5
TC visual inspection | good
Thickness Link0 (<650) | 618
Thickness Link1 (<650) | 617
Bow (long) Link0 (< 75) | 26
Bow (long) Link1 (< 75) | 28
Bow (across) Link0 (< 75) | 49
Bow (across) Link1 (< 75) | 58
Twist Link0 (<100) | -49
Twist Link1 (<100) | -44
Capacitance Vcc-GND (microF) at 1KHz (4.0<X<4.8) | 11.5
Capacitance Vdd-GND (microF) at 1KHz (4.0<X<4.8) | 11.47
Impedance HV-GND (kOhm) at 100Hz ( 30<X< 36) | 34.4
Impedance HV-GND (kOhm) at 1KHz ( 9<X< 11) | 10.1
Impedance HV-GND (kOhm) at 10KHz (5.7<X<7.7) | 6.75
Impedance HV-GND (kOhm) at 100KHz (5.0<X<7.1) | 6.08
Resistance R27 (Ohm) ( 99<X<101) | 100.7
Resistance R28 (Ohm) ( 99<X<101) | 100.2
Resistance R29 (Ohm) ( 99<X<101) | 100.2
Resistance R30 (Ohm) ( 99<X<101) | 101.1
Resistance TM1 (kOhm) at 25 deg.C. ( 9.8<X<10.2) | 10.03
Resistance TM2 (kOhm) at 25 deg.C. ( 9.8<X<10.2) | 10.2
Resistance ASIC die pad No.1-No.6 (milliOhm) (X<4) | 3.1
Resistance ASIC die pad No.7-No.12 (milliOhm) (X<4) | 3.7
LV leakage current for Icc (nA) at 10V (X<10) | 49
LV leakage current for Idd (nA) at 10V (X<10) | 49
HV leakge current (nA) at 500V (X<10) | 9

## ASIC_HYBRID (2)
atlasPartsId | 20220170200005
manufacturerSerialNo | <data>
eventDescription | ASIC_HYBRID
dateAssembly [dd/mm/yyyy] | 17/11/2001
location [instituteCode(DB)] | KEK
personInitial
problem [YES/NO] | NO
pass [YES/NO] | YES
comments
visualInspection | GOOD
number of Wirebond-Pull | 4
average [gr] | 9
max-min [gr] | 8.7
min [gr] | 1.7
typeASIC | ABCD3TA
ASIC m00 [siteID-lotID-waferID-Xnn-Ynn] | SCIPP-Z37277-W19-X06-Y03
s01 | SCIPP-Z37277-W19-X10-Y03
s02 | SCIPP-Z36459A-W07-X07-Y05 | <-replaced to
s03 | SCIPP-Z37277-W19-X07-Y01
s04 | SCIPP-Z37277-W19-X03-Y04
e05 | SCIPP-Z37277-W19-X13-Y05
m08 | SCIPP-Z37277-W19-X07-Y04
s09 | SCIPP-Z37277-W19-X15-Y06
s10 | SCIPP-Z37277-W19-X11-Y04
s11 | SCIPP-Z37277-W19-X10-Y04
s12 | SCIPP-Z37277-W19-X11-Y02
e13 | SCIPP-Z36459A-W07-X07-Y00
capCorrectionFactor m00 | 1.12 | (1)
s01 | 1.12 | (1)
s02 | 1.13 | (1)
s03 | 1.12 | (1)
s04 | 1.12 | (1)
e05 | 1.12 | (1)
m08 | 1.12 | (1)
s09 | 1.12 | (1)
s10 | 1.12 | (1)
s11 | 1.12 | (1)
s12 | 1.12 | (1)
e13 | 1.13 | (1)
JigID | H1
EOTITEP102 | 20220170131004
cureTemperature(xx) [C] | 50
cureDurationTime [hrs] | 2
dateWireBonding [dd/mm/yyyy] | 16/11/2001
location [institueCode(DB)]
nDefectiveChannels |  | (2)
defectiveChannels |  | (3)
comment
dateELECTRICAL-INITIAL-ASIC_HYBRID TEST | 17/11/2001
dateLONGTERM-ASIC_HYBRID | 22/11/2001
dateELECTRICAL-LT-ASIC_HYBRID | 22/11/2001
Notes:
(1) copy from ASIC datasheet
(2) number of defective channels
(3) n1,n2,n3,...(separated with comma)

## ASIC_HYBRID
atlasPartsId | 20220170200005
manufacturerSerialNo | <data>
eventDescription | ASIC_HYBRID
dateAssembly [dd/mm/yyyy] | 07/11/2001
location [instituteCode(DB)] | KEK
personInitial
problem [YES/NO] | YES
pass [YES/NO] | NO
comments | discard, chip to be replaced
visualInspection | GOOD
number of Wirebond-Pull | 4
average [gr] | 9
max-min [gr] | 8.7
min [gr] | 1.7
typeASIC | ABCD3TA
ASIC m00 [siteID-lotID-waferID-Xnn-Ynn] | SCIPP-Z37277-W19-X06-Y03
s01 | SCIPP-Z37277-W19-X10-Y03
s02 | SCIPP-Z37277-W19-X11-Y06 | ->replace
s03 | SCIPP-Z37277-W19-X07-Y01
s04 | SCIPP-Z37277-W19-X03-Y04
e05 | SCIPP-Z37277-W19-X13-Y05
m08 | SCIPP-Z37277-W19-X07-Y04
s09 | SCIPP-Z37277-W19-X15-Y06
s10 | SCIPP-Z37277-W19-X11-Y04
s11 | SCIPP-Z37277-W19-X10-Y04
s12 | SCIPP-Z37277-W19-X11-Y02
e13 | SCIPP-Z36459A-W07-X07-Y00
capCorrectionFactor m00 | 1.12 | (1)
s01 | 1.12 | (1)
s02 | 1.12 | (1)
s03 | 1.12 | (1)
s04 | 1.12 | (1)
e05 | 1.12 | (1)
m08 | 1.12 | (1)
s09 | 1.12 | (1)
s10 | 1.12 | (1)
s11 | 1.12 | (1)
s12 | 1.12 | (1)
e13 | 1.13 | (1)
JigID | H1
EOTITEP102 | 20220170131004
cureTemperature(xx) [C] | 50
cureDurationTime [hrs] | 2
dateWireBonding [dd/mm/yyyy] | 07/11/2001
location [institueCode(DB)]
nDefectiveChannels |  | (2)
defectiveChannels |  | (3)
comment
dateELECTRICAL-INITIAL-ASIC_HYBRID TEST
dateLONGTERM-ASIC_HYBRID
dateELECTRICAL-LT-ASIC_HYBRID
Notes:
(1) copy from ASIC datasheet
(2) number of defective channels
(3) n1,n2,n3,...(separated with comma)

## ELECTRICAL-INITIAL-AH
atlasPartsId | 20220170200005
eventDescription | ELECTRICAL-INITIAL-ASIC_HYBRID |  | : X=INITIAL, TC, LT, ANNEAL, ...
date [dd/mm/yyyy] | 16/11/2001
location [instituteCode(DB)] | KEK
personInitial | ta
problem [YES/NO] | no
pass [YES/NO] | yes
comments | replacing chip2
#
# | http://hepwww.rl.ac.uk/atlas-sct/documents/dev/Electrical_Tests_301.doc
# | N.B. Some revisions have been incorporated here, but not yet in the reference,
# | namely that CONFIGURATION will be stored as an ASCII file and
# | and that the RUN field will be used to store both RUN and SCAN number.
#
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | FullBypassTest
%TEST_DATE | 16/11/2001
%TEST_TIME | 23:41:10
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4612
%SCAN | 1
#
%HOST | ATLASPC9
%T0 | 36.0 | deg. C
%T1 | 37.0 | deg. C
#
#NConfigs 37
#Vdd = 4.2 - 3.3
#
%M0->S1: | 3.3 V | (Minimal tested)
%M0->S2: | 3.3 V | (Minimal tested)
%S1->S2: | 3.3 V | (Minimal tested)
%S1->S3: | 3.3 V | (Minimal tested)
%S2->S3: | 3.3 V | (Minimal tested)
%S2->S4: | 3.3 V | (Minimal tested)
%S3->S4: | 3.3 V | (Minimal tested)
%S3->E5: | 3.3 V | (Minimal tested)
%S4->E5: | 3.3 V | (Minimal tested)
%E5->S9: | 3.3 V | (Minimal tested)
%M8->S9: | 3.3 V | (Minimal tested)
%M8->S10: | 3.3 V | (Minimal tested)
%S9->S10: | 3.3 V | (Minimal tested)
%S9->S11: | 3.3 V | (Minimal tested)
%S10->S11: | 3.3 V | (Minimal tested)
%S10->S12: | 3.3 V | (Minimal tested)
%S11->S12: | 3.3 V | (Minimal tested)
%S11->E13: | 3.3 V | (Minimal tested)
%S12->E13: | 3.3 V | (Minimal tested)
%E13->S1: | 3.3 V | (Minimal tested)
#
%DEFECTS
#name | bypass_link | Vdd_min
#No defects found!
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | PipelineTest
%TEST_DATE | 16/11/2001
%TEST_TIME | 23:41:21
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4612
%SCAN | 11
#
%HOST | ATLASPC9
%T0 | 34.0 | deg. C
%T1 | 36.0 | deg. C
#
#
%00_GOOD | 128
%01_GOOD | 128
%02_GOOD | 128
%03_GOOD | 128
%04_GOOD | 128
%05_GOOD | 128
%06_GOOD | 128
%07_GOOD | 128
%08_GOOD | 128
%09_GOOD | 128
%10_GOOD | 128
%11_GOOD | 128
#
%DEFECTS
#name | chan | cell
#No defects found!
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | StrobeDelay
%TEST_DATE | 16/11/2001
%TEST_TIME | 23:44:57
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4613
%SCAN | 2
#
%HOST | ATLASPC9
%T0 | 37.0 | deg. C
%T1 | 38.0 | deg. C
#
#
#parameter | value
%00_DELAY | 12
%01_DELAY | 14
%02_DELAY | 14
%03_DELAY | 14
%04_DELAY | 14
%05_DELAY | 14
%06_DELAY | 13
%07_DELAY | 13
%08_DELAY | 14
%09_DELAY | 14
%10_DELAY | 14
%11_DELAY | 16
#
%DEFECTS
#name | chip
#No defects found!
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | ResponseCurve
%TEST_DATE | 16/11/2001
%TEST_TIME | 23:49:53
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4614
%SCAN | 1
#
%HOST | ATLASPC9
%T0 | 37.0 | deg. C
%T1 | 37.0 | deg. C
#
#
%N_SCANS | 3
%Q_00 | 1.50 fC
%Q_01 | 2.00 fC
%Q_02 | 2.50 fC
%RC_FIT | 4 | (linear fit)
#
%ModuleResponseCurvePlot
# |  | func | p0 | p1 | p2
Chip 0: | 4 | 18.29 | 54.13 | 0.00
Chip 1: | 4 | 16.98 | 55.29 | 0.00
Chip 2: | 4 | 12.06 | 54.17 | 0.00
Chip 3: | 4 | 16.26 | 55.76 | 0.00
Chip 4: | 4 | 15.24 | 57.14 | 0.00
Chip 5: | 4 | 18.51 | 54.05 | 0.00
Chip 6: | 4 | 16.73 | 54.17 | 0.00
Chip 7: | 4 | 13.44 | 55.94 | 0.00
Chip 8: | 4 | 14.56 | 54.76 | 0.00
Chip 9: | 4 | 13.61 | 54.38 | 0.00
Chip 10: | 4 | 12.32 | 54.93 | 0.00
Chip 11: | 4 | 16.32 | 53.53 | 0.00
#
%ModuleGainOffsetNoise
# at 2.00fC
# |  | vt50 | rms | gain | rms | offset | rms | outnse | innse | rms | comment
Chip 0: | 126.2 | 7.48 | 54.1 | 0.95 | 19.0 | 6.85 | 5.66 | 654 | 20 | unbonded
Chip 1: | 127.7 | 7.50 | 55.3 | 0.97 | 17.1 | 7.83 | 5.69 | 643 | 20 | unbonded
Chip 2: | 121.0 | 6.66 | 54.1 | 1.12 | 12.4 | 6.01 | 5.27 | 608 | 20 | unbonded
Chip 3: | 128.1 | 7.40 | 55.8 | 1.05 | 16.8 | 6.57 | 5.67 | 636 | 21 | unbonded
Chip 4: | 129.9 | 7.78 | 57.1 | 1.02 | 15.9 | 6.86 | 5.54 | 606 | 16 | unbonded
Chip 5: | 126.6 | 8.19 | 54.0 | 0.89 | 18.9 | 7.85 | 5.56 | 643 | 22 | unbonded
Chip 6: | 125.1 | 7.56 | 54.2 | 0.84 | 17.2 | 7.29 | 5.59 | 645 | 18 | unbonded
Chip 7: | 125.6 | 6.55 | 55.9 | 0.95 | 13.5 | 6.73 | 5.69 | 636 | 18 | unbonded
Chip 8: | 124.0 | 7.69 | 54.8 | 0.86 | 15.0 | 7.08 | 5.62 | 641 | 20 | unbonded
Chip 9: | 122.7 | 8.39 | 54.4 | 0.81 | 14.7 | 7.48 | 5.53 | 635 | 20 | unbonded
Chip 10: | 122.2 | 7.86 | 54.9 | 1.00 | 12.9 | 7.30 | 5.55 | 632 | 16 | unbonded
Chip 11: | 123.2 | 7.51 | 53.5 | 1.03 | 16.6 | 7.39 | 5.54 | 647 | 21 | unbonded
#
#BadChannelSummary - not for the database
# at 2.00fC
# |  | lost | dodgy | dead | stuck | ineff | unbon | lo_gn | hi_gn | lo_off | hi_off | partbon | hi_nse
#Chip 0: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 1: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 2: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 3: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 4: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 5: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 6: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 7: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 8: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 9: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 10: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 11: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Link 0: | 768 | 0 | 0 | 0 | 0 | 768 | 0 | 0 | 0 | 0 | 0 | 0
#Link 1: | 768 | 0 | 0 | 0 | 0 | 768 | 0 | 0 | 0 | 0 | 0 | 0
#Link 2: | 1536 | 0 | 0 | 0 | 0 | 1536 | 0 | 0 | 0 | 0 | 0 | 0
#
%DEFECTS
#name | channel | value
#No defects found!
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 17/11/2001
%TEST_TIME | 00:05:51
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | NO
%PROBLEM | NO
#
%RUN | 4615
%SCAN | 1
#
%HOST | ATLASPC9
%T0 | 38.0 | deg. C
%T1 | 38.0 | deg. C
#
%TRIM_TYPE | 0
%TRIM_Q | 10.0
%N_SCANS | 16
%T_00 | 0
%T_01 | 1
%T_02 | 2
%T_03 | 3
%T_04 | 4
%T_05 | 5
%T_06 | 6
%T_07 | 7
%T_08 | 8
%T_09 | 9
%T_10 | 10
%T_11 | 11
%T_12 | 12
%T_13 | 13
%T_14 | 14
%T_15 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005hb220s30_tr0_20011117.trim
%MASK_FILE | D:\sctvar\results\0170200005hb220s30_tr0_20011117.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 0 | 92.5 | 128 | 92.3 | 1.22 | -18.1 | 2.0 | 3.87 | 0.18
Chip 1: | 0 | 92.5 | 128 | 92.2 | 1.11 | -19.8 | 2.5 | 3.53 | 0.17
Chip 2: | 0 | 92.5 | 128 | 91.6 | 1.01 | -22.4 | 2.8 | 2.99 | 0.26
Chip 3: | 0 | 92.5 | 128 | 91.9 | 1.25 | -18.6 | 1.9 | 3.79 | 0.16
Chip 4: | 0 | 92.5 | 128 | 91.7 | 1.17 | -20.4 | 2.0 | 3.51 | 0.19
Chip 5: | 0 | 92.5 | 128 | 92.2 | 1.04 | -20.7 | 2.5 | 3.38 | 0.11
Chip 6: | 0 | 92.5 | 128 | 92.1 | 1.18 | -19.4 | 2.2 | 3.56 | 0.16
Chip 7: | 0 | 92.5 | 128 | 91.8 | 1.04 | -19.7 | 2.0 | 3.51 | 0.18
Chip 8: | 0 | 92.5 | 128 | 91.7 | 1.05 | -19.7 | 2.4 | 3.44 | 0.15
Chip 9: | 0 | 92.5 | 128 | 91.9 | 1.07 | -20.1 | 2.5 | 3.34 | 0.14
Chip 10: | 0 | 92.5 | 128 | 92.0 | 1.08 | -19.1 | 2.3 | 3.44 | 0.13
Chip 11: | 0 | 92.5 | 127 | 92.1 | 1.28 | -16.7 | 1.8 | 4.11 | 0.36
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -11.32
%TR_STEP | 127 | 4.42
%TR_STEP | 165 | 4.61
%TR_OFFSET | 452 | -11.55
%TR_STEP | 511 | 4.36
%TR_OFFSET | 608 | -13.81
%TR_STEP | 639 | 4.14
%TR_OFFSET | 691 | -13.01
%TR_OFFSET | 734 | -13.11
%TR_OFFSET | 171 | -26.35
%TR_OFFSET | 210 | -26.25
%TR_STEP | 251 | 4.14
%TR_STEP | 255 | 4.17
%TR_OFFSET | 585 | -26.25
%TR_OFFSET | 750 | -23.73
%TR_STEP | 765 | 5.25
#16 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 17/11/2001
%TEST_TIME | 00:05:51
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 4615
%SCAN | 17
#
%HOST | ATLASPC9
%T0 | 38.0 | deg. C
%T1 | 38.0 | deg. C
#
%TRIM_TYPE | 1
%TRIM_Q | 10.0
%N_SCANS | 4
%T_00 | 3
%T_01 | 7
%T_02 | 11
%T_03 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005hb220s30_tr1_20011117.trim
%MASK_FILE | D:\sctvar\results\0170200005hb220s30_tr1_20011117.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 1 | 115.0 | 128 | 115.2 | 2.18 | -9.9 | 1.0 | 7.16 | 0.29
Chip 1: | 1 | 115.0 | 128 | 115.3 | 1.97 | -10.5 | 1.5 | 6.70 | 0.41
Chip 2: | 1 | 115.0 | 128 | 114.8 | 1.82 | -10.7 | 1.3 | 6.16 | 0.42
Chip 3: | 1 | 115.0 | 128 | 114.9 | 2.06 | -9.9 | 1.1 | 7.16 | 0.27
Chip 4: | 1 | 115.0 | 128 | 114.6 | 1.91 | -10.6 | 1.2 | 6.79 | 0.31
Chip 5: | 1 | 115.0 | 128 | 115.2 | 1.86 | -11.0 | 1.4 | 6.44 | 0.21
Chip 6: | 1 | 115.0 | 128 | 115.1 | 1.91 | -10.2 | 1.2 | 6.79 | 0.27
Chip 7: | 1 | 115.0 | 128 | 115.3 | 1.97 | -9.9 | 1.1 | 6.94 | 0.31
Chip 8: | 1 | 115.0 | 128 | 115.2 | 2.08 | -10.1 | 1.2 | 6.69 | 0.27
Chip 9: | 1 | 115.0 | 128 | 115.1 | 1.81 | -10.3 | 1.4 | 6.50 | 0.25
Chip 10: | 1 | 115.0 | 128 | 115.1 | 1.96 | -9.8 | 1.2 | 6.73 | 0.25
Chip 11: | 1 | 115.0 | 128 | 114.7 | 2.25 | -8.9 | 1.0 | 7.70 | 0.57
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -6.03
%TR_STEP | 123 | 8.06
%TR_STEP | 127 | 8.04
%TR_STEP | 165 | 10.53
%TR_OFFSET | 452 | -6.08
%TR_STEP | 511 | 8.06
%TR_STEP | 639 | 7.85
%TR_OFFSET | 691 | -6.61
%TR_OFFSET | 734 | -6.52
%TR_OFFSET | 171 | -13.38
%TR_OFFSET | 210 | -13.83
%TR_STEP | 251 | 7.90
%TR_STEP | 255 | 8.00
%TR_OFFSET | 585 | -13.77
%TR_STEP | 623 | 7.60
%TR_STEP | 627 | 7.49
%TR_OFFSET | 750 | -12.49
%TR_STEP | 765 | 9.52
#18 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 17/11/2001
%TEST_TIME | 00:05:51
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 4615
%SCAN | 21
#
%HOST | ATLASPC9
%T0 | 38.0 | deg. C
%T1 | 38.0 | deg. C
#
%TRIM_TYPE | 2
%TRIM_Q | 10.0
%N_SCANS | 4
%T_00 | 3
%T_01 | 7
%T_02 | 11
%T_03 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005hb220s30_tr2_20011117.trim
%MASK_FILE | D:\sctvar\results\0170200005hb220s30_tr2_20011117.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 2 | 140.0 | 128 | 140.0 | 2.98 | -6.6 | 0.6 | 10.63 | 0.45
Chip 1: | 2 | 140.0 | 128 | 139.6 | 2.94 | -7.0 | 1.0 | 9.95 | 0.61
Chip 2: | 2 | 140.0 | 128 | 140.1 | 2.66 | -7.1 | 0.8 | 9.27 | 0.57
Chip 3: | 2 | 140.0 | 128 | 139.9 | 2.91 | -6.6 | 0.7 | 10.66 | 0.39
Chip 4: | 2 | 140.0 | 128 | 139.9 | 2.86 | -7.1 | 0.8 | 10.09 | 0.43
Chip 5: | 2 | 140.0 | 128 | 140.4 | 2.81 | -7.4 | 0.8 | 9.60 | 0.33
Chip 6: | 2 | 140.0 | 128 | 140.5 | 2.98 | -6.9 | 0.7 | 10.10 | 0.41
Chip 7: | 2 | 140.0 | 128 | 140.2 | 3.00 | -6.5 | 0.8 | 10.44 | 0.43
Chip 8: | 2 | 140.0 | 128 | 140.2 | 2.81 | -6.8 | 0.8 | 9.90 | 0.39
Chip 9: | 2 | 140.0 | 128 | 140.1 | 2.78 | -6.9 | 0.9 | 9.65 | 0.37
Chip 10: | 2 | 140.0 | 128 | 139.8 | 2.84 | -6.6 | 0.8 | 9.96 | 0.37
Chip 11: | 2 | 140.0 | 128 | 140.2 | 3.22 | -6.2 | 0.7 | 11.15 | 0.76
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -4.27
%TR_STEP | 165 | 15.60
%TR_OFFSET | 452 | -4.12
%TR_STEP | 511 | 11.98
%TR_STEP | 639 | 11.51
%TR_OFFSET | 691 | -4.66
%TR_OFFSET | 734 | -4.71
%TR_OFFSET | 171 | -9.15
%TR_OFFSET | 210 | -9.07
%TR_OFFSET | 585 | -9.36
%TR_STEP | 623 | 11.10
%TR_OFFSET | 750 | -10.59
%TR_STEP | 765 | 13.43
%TR_STEP | 766 | 13.59
#14 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 17/11/2001
%TEST_TIME | 00:05:51
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 4615
%SCAN | 25
#
%HOST | ATLASPC9
%T0 | 38.0 | deg. C
%T1 | 38.0 | deg. C
#
%TRIM_TYPE | 3
%TRIM_Q | 10.0
%N_SCANS | 4
%T_00 | 3
%T_01 | 7
%T_02 | 11
%T_03 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005hb220s30_tr3_20011117.trim
%MASK_FILE | D:\sctvar\results\0170200005hb220s30_tr3_20011117.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 3 | 160.0 | 128 | 160.6 | 4.14 | -4.9 | 0.5 | 14.10 | 0.56
Chip 1: | 3 | 160.0 | 128 | 159.3 | 3.60 | -5.2 | 0.8 | 13.28 | 0.80
Chip 2: | 3 | 160.0 | 128 | 159.8 | 3.65 | -5.3 | 0.6 | 12.36 | 0.71
Chip 3: | 3 | 160.0 | 128 | 159.9 | 4.03 | -4.9 | 0.5 | 14.10 | 0.50
Chip 4: | 3 | 160.0 | 128 | 160.1 | 3.90 | -5.4 | 0.6 | 13.30 | 0.57
Chip 5: | 3 | 160.0 | 128 | 159.7 | 3.81 | -5.4 | 0.6 | 12.90 | 0.43
Chip 6: | 3 | 160.0 | 128 | 160.3 | 4.22 | -5.1 | 0.5 | 13.46 | 0.53
Chip 7: | 3 | 160.0 | 128 | 159.8 | 3.97 | -5.0 | 0.6 | 13.77 | 0.55
Chip 8: | 3 | 160.0 | 128 | 160.1 | 3.97 | -5.1 | 0.6 | 13.16 | 0.51
Chip 9: | 3 | 160.0 | 128 | 159.8 | 3.70 | -5.2 | 0.7 | 12.83 | 0.48
Chip 10: | 3 | 160.0 | 128 | 159.6 | 3.66 | -4.9 | 0.6 | 13.23 | 0.49
Chip 11: | 3 | 160.0 | 128 | 160.1 | 4.48 | -4.6 | 0.5 | 14.75 | 0.89
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -3.17
%TR_STEP | 165 | 20.75
%TR_STEP | 258 | 14.57
%TR_OFFSET | 452 | -3.06
%TR_STEP | 511 | 15.71
%TR_STEP | 639 | 15.10
%TR_OFFSET | 691 | -3.53
%TR_OFFSET | 734 | -3.45
%TR_OFFSET | 171 | -6.65
%TR_OFFSET | 210 | -6.81
%TR_STEP | 255 | 15.43
%TR_STEP | 351 | 14.71
%TR_OFFSET | 585 | -6.84
%TR_STEP | 627 | 14.71
%TR_OFFSET | 750 | -7.66
%TR_STEP | 765 | 17.61
#16 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 17/11/2001
%TEST_TIME | 00:05:51
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 4615
%SCAN | 1
#
%HOST | ATLASPC9
%T0 | 38.0 | deg. C
%T1 | 38.0 | deg. C
#
%TRIM_TYPE | -1
%TRIM_Q | 10.0
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005hb220s30_tr-1_20011117.trim
%MASK_FILE | D:\sctvar\results\0170200005hb220s30_tr-1_20011117.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 0 | 95.0 | 128 | 95.1 | 1.11 | -18.1 | 2.0 | 3.87 | 0.18
Chip 1: | 0 | 95.0 | 128 | 95.1 | 1.03 | -19.8 | 2.5 | 3.53 | 0.17
Chip 2: | 1 | 95.0 | 128 | 95.3 | 1.69 | -10.7 | 1.3 | 6.16 | 0.42
Chip 3: | 0 | 95.0 | 128 | 94.8 | 1.05 | -18.6 | 1.9 | 3.79 | 0.16
Chip 4: | 0 | 95.0 | 128 | 95.0 | 0.93 | -20.4 | 2.0 | 3.51 | 0.19
Chip 5: | 0 | 95.0 | 128 | 95.1 | 0.95 | -20.7 | 2.5 | 3.38 | 0.11
Chip 6: | 0 | 95.0 | 128 | 94.9 | 1.11 | -19.4 | 2.2 | 3.56 | 0.16
Chip 7: | 0 | 95.0 | 128 | 95.1 | 1.01 | -19.7 | 2.0 | 3.51 | 0.18
Chip 8: | 0 | 95.0 | 128 | 95.0 | 0.87 | -19.7 | 2.4 | 3.44 | 0.15
Chip 9: | 0 | 95.0 | 128 | 95.0 | 0.97 | -20.1 | 2.5 | 3.34 | 0.14
Chip 10: | 0 | 95.0 | 128 | 95.0 | 1.05 | -19.1 | 2.3 | 3.44 | 0.13
Chip 11: | 1 | 95.0 | 128 | 95.2 | 2.42 | -8.9 | 1.0 | 7.70 | 0.57
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -11.32
%TR_STEP | 127 | 4.42
%TR_STEP | 165 | 4.61
%TR_OFFSET | 452 | -11.55
%TR_STEP | 511 | 4.36
%TR_OFFSET | 608 | -13.81
%TR_STEP | 639 | 4.14
%TR_OFFSET | 691 | -13.01
%TR_OFFSET | 734 | -13.11
%TR_OFFSET | 171 | -26.35
%TR_OFFSET | 210 | -26.25
%TR_STEP | 251 | 4.14
%TR_STEP | 255 | 4.17
%TR_OFFSET | 585 | -26.25
%TR_OFFSET | 750 | -12.49
%TR_STEP | 765 | 9.52
#16 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30q1t092r0ini
%TEST_NAME | ResponseCurve
%TEST_DATE | 17/11/2001
%TEST_TIME | 00:31:21
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4618
%SCAN | 1
#
%HOST | ATLASPC9
%T0 | 36.0 | deg. C
%T1 | 36.0 | deg. C
#
#
%N_SCANS | 10
%Q_00 | 0.50 fC
%Q_01 | 0.75 fC
%Q_02 | 1.00 fC
%Q_03 | 1.25 fC
%Q_04 | 1.50 fC
%Q_05 | 2.00 fC
%Q_06 | 3.00 fC
%Q_07 | 4.00 fC
%Q_08 | 6.00 fC
%Q_09 | 8.00 fC
%RC_FIT | 3 | (exponential fit)
#
%ModuleResponseCurvePlot
# |  | func | p0 | p1 | p2
Chip 0: | 3 | 1435.90 | 6.25 | -683.36
Chip 1: | 3 | 1370.09 | 5.89 | -650.84
Chip 2: | 3 | 1233.21 | 5.54 | -580.18
Chip 3: | 3 | 1380.90 | 5.92 | -656.50
Chip 4: | 3 | 1389.10 | 5.90 | -660.91
Chip 5: | 3 | 1302.82 | 5.68 | -616.40
Chip 6: | 3 | 1362.72 | 6.04 | -645.39
Chip 7: | 3 | 1402.87 | 6.14 | -665.93
Chip 8: | 3 | 1334.26 | 5.86 | -631.89
Chip 9: | 3 | 1332.96 | 5.96 | -630.02
Chip 10: | 3 | 1315.28 | 5.89 | -620.55
Chip 11: | 3 | 1540.62 | 6.92 | -733.82
#
%ModuleGainOffsetNoise
# at 2.00fC
# |  | vt50 | rms | gain | rms | offset | rms | outnse | innse | rms | comment
Chip 0: | 147.7 | 1.46 | 56.0 | 0.97 | 34.6 | 1.68 | 5.46 | 610 | 19 | unbonded
Chip 1: | 149.2 | 1.49 | 56.5 | 0.99 | 34.2 | 1.61 | 5.43 | 600 | 17 | unbonded
Chip 2: | 147.2 | 1.33 | 53.7 | 1.11 | 36.6 | 1.72 | 5.23 | 608 | 20 | unbonded
Chip 3: | 149.4 | 1.66 | 56.6 | 1.12 | 34.1 | 1.70 | 5.41 | 598 | 20 | unbonded
Chip 4: | 150.7 | 1.49 | 57.1 | 1.12 | 33.9 | 1.68 | 5.29 | 579 | 18 | unbonded
Chip 5: | 148.1 | 1.39 | 55.4 | 1.06 | 35.3 | 1.63 | 5.26 | 594 | 20 | unbonded
Chip 6: | 147.8 | 1.42 | 54.9 | 0.92 | 36.0 | 1.61 | 5.34 | 607 | 22 | unbonded
Chip 7: | 149.1 | 1.35 | 55.5 | 0.92 | 35.7 | 1.45 | 5.51 | 620 | 19 | unbonded
Chip 8: | 148.0 | 1.38 | 55.3 | 0.99 | 35.4 | 1.53 | 5.47 | 619 | 20 | unbonded
Chip 9: | 147.8 | 1.20 | 54.4 | 0.90 | 36.5 | 1.52 | 5.33 | 612 | 18 | unbonded
Chip 10: | 148.3 | 1.35 | 54.3 | 1.05 | 37.1 | 1.60 | 5.39 | 621 | 17 | unbonded
Chip 11: | 146.8 | 1.61 | 54.5 | 1.20 | 36.5 | 1.87 | 5.42 | 621 | 21 | unbonded
#
#BadChannelSummary - not for the database
# at 2.00fC
# |  | lost | dodgy | dead | stuck | ineff | unbon | lo_gn | hi_gn | lo_off | hi_off | partbon | hi_nse
#Chip 0: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 1: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 2: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 3: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 4: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 5: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 6: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 7: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 8: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 9: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 10: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 11: | 128 | 0 | 1 | 0 | 0 | 127 | 0 | 0 | 0 | 0 | 0 | 0
#Link 0: | 768 | 0 | 0 | 0 | 0 | 768 | 0 | 0 | 0 | 0 | 0 | 0
#Link 1: | 768 | 0 | 1 | 0 | 0 | 767 | 0 | 0 | 0 | 0 | 0 | 0
#Link 2: | 1536 | 0 | 1 | 0 | 0 | 1535 | 0 | 0 | 0 | 0 | 0 | 0
#
%DEFECTS
#name | channel | value
%DEAD | 1518
#1 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30q1t092r0ini
%TEST_NAME | Timewalk
%TEST_DATE | 17/11/2001
%TEST_TIME | 00:45:42
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4619
%SCAN | 2
#
%HOST | ATLASPC9
%T0 | 38.0 | deg. C
%T1 | 38.0 | deg. C
#
#
# run scan charge
# 4619 2 1.25
# 4619 3 1.50
# 4619 4 1.75
# 4619 5 2.00
# 4619 6 3.00
# 4619 7 4.00
# 4619 8 5.00
# 4619 9 6.00
# 4619 10 8.00
# 4619 11 10.00
#
%00_TW | 9.5
%01_TW | 9.6
%02_TW | 9.9
%03_TW | 9.9
%04_TW | 9.9
%05_TW | 9.9
%06_TW | 9.8
%07_TW | 10.3
%08_TW | 10.2
%09_TW | 10.1
%10_TW | 9.9
%11_TW | 9.1
#
%00_TCAL | 28.6
%01_TCAL | 28.5
%02_TCAL | 29.8
%03_TCAL | 28.8
%04_TCAL | 28.7
%05_TCAL | 28.6
%06_TCAL | 28.6
%07_TCAL | 28.3
%08_TCAL | 28.5
%09_TCAL | 28.1
%10_TCAL | 28.1
%11_TCAL | 28.7
#
%DEFECTS
#name | chip | timewalk
#No defects found!

## LONGTERM-AH
atlasPartsId | 20220170200005
eventDescription | LONGTERM-ASIC_HYBRID | : X=ASIC_HYBRID, MODULE
dateEntry [dd/mm/yyyy] | 22/11/2001
location [instituteCode(DB)] | KEK
personInitial | ta
problem [YES/NO] | no
pass [YES/NO] | yes
comments | no
dateStart [dd/mm/yyyy] | 18/11/2001
Long term test length (hours) | 104
Long term test maximum T1(xx.x) (C) | 37
Long term test minimum T1(xx.x) (C) | 33
Long term test maximum T2(xx.x) (C) | 38
Long term test minimum T2(xx.x) (C) | 34
Long term test maximum Icc (mA) | 920
Long term test minimum Icc (mA) | 890
Long term test maximum Idd (mA) | 550
Long term test minimum Idd (mA) | 420
Long term test time of first failure | no
Long term test failed ASICs | no | chip1,chip2,... (separated with comma)
RAWDATA
Time [hr] | Temp1 [C] | Temp2 [C] | Icc [mA] | Idd [mA] | every 1hr
0 | 34 | 35 | 920 | 490
1 | 37 | 38 | 910 | 490
2 | 37 | 37 | 920 | 490
3 | 37 | 38 | 910 | 490
4 | 37 | 37 | 920 | 490
5 | 37 | 38 | 910 | 490
6 | 37 | 37 | 910 | 490
7 | 37 | 38 | 910 | 490
8 | 37 | 37 | 910 | 490
9 | 37 | 37 | 910 | 490
10 | 37 | 37 | 910 | 490
11 | 37 | 37 | 910 | 490
12 | 37 | 37 | 910 | 490
13 | 36 | 37 | 910 | 490
14 | 37 | 37 | 910 | 490
15 | 37 | 38 | 910 | 490
16 | 37 | 34 | 910 | 490
17 | 37 | 37 | 920 | 470
18 | 37 | 37 | 920 | 490
19 | 37 | 37 | 910 | 490
20 | 37 | 37 | 910 | 490
21 | 37 | 38 | 910 | 490
22 | 37 | 37 | 910 | 490
23 | 37 | 37 | 910 | 490
24 | 37 | 38 | 910 | 490
25 | 37 | 37 | 920 | 490
26 | 37 | 38 | 910 | 490
27 | 37 | 38 | 910 | 490
28 | 37 | 37 | 910 | 490
29 | 37 | 35 | 910 | 490
30 | 37 | 36 | 910 | 490
31 | 37 | 37 | 910 | 490
32 | 36 | 37 | 910 | 490
33 | 37 | 37 | 910 | 480
34 | 37 | 37 | 910 | 490
35 | 37 | 38 | 910 | 490
36 | 37 | 37 | 910 | 490
37 | 37 | 37 | 910 | 490
38 | 36 | 37 | 910 | 490
39 | 33 | 37 | 910 | 490
40 | 37 | 37 | 910 | 490
41 | 37 | 37 | 920 | 490
42 | 37 | 37 | 910 | 490
43 | 37 | 34 | 910 | 490
44 | 34 | 37 | 910 | 490
45 | 37 | 38 | 910 | 490
46 | 37 | 37 | 910 | 490
47 | 37 | 37 | 910 | 490
48 | 37 | 37 | 910 | 490
49 | 37 | 37 | 920 | 470
50 | 37 | 37 | 910 | 490
51 | 37 | 37 | 910 | 490
52 | 37 | 37 | 910 | 490
53 | 36 | 37 | 910 | 490
54 | 37 | 37 | 910 | 490
55 | 37 | 37 | 920 | 480
56 | 37 | 38 | 910 | 490
57 | 37 | 37 | 910 | 490
58 | 37 | 37 | 910 | 490
59 | 37 | 37 | 910 | 490
60 | 37 | 37 | 910 | 490
61 | 37 | 37 | 910 | 490
62 | 37 | 34 | 910 | 490
63 | 37 | 38 | 910 | 490
64 | 36 | 37 | 920 | 450
65 | 37 | 37 | 920 | 480
66 | 37 | 37 | 920 | 480
67 | 37 | 38 | 920 | 480
68 | 37 | 38 | 910 | 490
69 | 37 | 37 | 920 | 420
70 | 36 | 37 | 920 | 420
71 | 37 | 38 | 920 | 480
72 | 37 | 38 | 890 | 550
73 | 36 | 36 | 920 | 480
74 | 37 | 38 | 910 | 490
75 | 37 | 37 | 910 | 490
76 | 37 | 37 | 910 | 490
77 | 37 | 38 | 910 | 480
78 | 37 | 38 | 910 | 490
79 | 37 | 37 | 910 | 480
80 | 37 | 38 | 910 | 480
81 | 37 | 38 | 910 | 480
82 | 36 | 37 | 910 | 480
83 | 37 | 38 | 910 | 500
84 | 37 | 37 | 910 | 480
85 | 37 | 37 | 920 | 490
86 | 37 | 38 | 910 | 490
87 | 37 | 38 | 910 | 490
88 | 37 | 38 | 910 | 480
89 | 37 | 38 | 910 | 490
90 | 37 | 38 | 910 | 490
91 | 37 | 38 | 920 | 500
92 | 37 | 38 | 920 | 480
93 | 37 | 35 | 910 | 490
94 | 37 | 38 | 910 | 490
95 | 37 | 38 | 910 | 490
96 | 37 | 37 | 920 | 480
97 | 37 | 38 | 910 | 480
98 | 37 | 38 | 920 | 490
99 | 37 | 38 | 910 | 490
100 | 37 | 38 | 910 | 490
101 | 37 | 38 | 910 | 490
102 | 37 | 38 | 910 | 510
103 | 36 | 38 | 920 | 490
104 | 37 | 38 | 910 | 490

## ELECTRICAL-LT-AH
atlasPartsId | 20220170200005
eventDescription | ELECTRICAL-LT-ASIC_HYBRID |  | : X=INITIAL, TC, LT, ANNEAL, ...
date [dd/mm/yyyy] | 22/11/2001
location [instituteCode(DB)] | KEK
personInitial | ta
problem [YES/NO] | no
pass [YES/NO] | yes
comments | no
#
# | http://hepwww.rl.ac.uk/atlas-sct/documents/dev/Electrical_Tests_301.doc
# | N.B. Some revisions have been incorporated here, but not yet in the reference,
# | namely that CONFIGURATION will be stored as an ASCII file and
# | and that the RUN field will be used to store both RUN and SCAN number.
#
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | FullBypassTest
%TEST_DATE | 22/11/2001
%TEST_TIME | 09:53:17
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4658
%SCAN | 1
#
%HOST | ATLASPC9
%T0 | 33.0 | deg. C
%T1 | 33.0 | deg. C
#
#NConfigs 37
#Vdd = 4.2 - 3.3
#
%M0->S1: | 3.3 V | (Minimal tested)
%M0->S2: | 3.3 V | (Minimal tested)
%S1->S2: | 3.3 V | (Minimal tested)
%S1->S3: | 3.3 V | (Minimal tested)
%S2->S3: | 3.3 V | (Minimal tested)
%S2->S4: | 3.3 V | (Minimal tested)
%S3->S4: | 3.3 V | (Minimal tested)
%S3->E5: | 3.3 V | (Minimal tested)
%S4->E5: | 3.3 V | (Minimal tested)
%E5->S9: | 3.3 V | (Minimal tested)
%M8->S9: | 3.3 V | (Minimal tested)
%M8->S10: | 3.3 V | (Minimal tested)
%S9->S10: | 3.3 V | (Minimal tested)
%S9->S11: | 3.3 V | (Minimal tested)
%S10->S11: | 3.3 V | (Minimal tested)
%S10->S12: | 3.3 V | (Minimal tested)
%S11->S12: | 3.3 V | (Minimal tested)
%S11->E13: | 3.3 V | (Minimal tested)
%S12->E13: | 3.3 V | (Minimal tested)
%E13->S1: | 3.3 V | (Minimal tested)
#
%DEFECTS
#name | bypass_link | Vdd_min
#No defects found!
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | PipelineTest
%TEST_DATE | 22/11/2001
%TEST_TIME | 09:53:28
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4658
%SCAN | 11
#
%HOST | ATLASPC9
%T0 | 33.0 | deg. C
%T1 | 33.0 | deg. C
#
#
%00_GOOD | 128
%01_GOOD | 128
%02_GOOD | 128
%03_GOOD | 128
%04_GOOD | 128
%05_GOOD | 128
%06_GOOD | 128
%07_GOOD | 128
%08_GOOD | 128
%09_GOOD | 128
%10_GOOD | 128
%11_GOOD | 128
#
%DEFECTS
#name | chan | cell
#No defects found!
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | StrobeDelay
%TEST_DATE | 22/11/2001
%TEST_TIME | 09:58:13
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4659
%SCAN | 2
#
%HOST | ATLASPC9
%T0 | 32.0 | deg. C
%T1 | 32.0 | deg. C
#
#
#parameter | value
%00_DELAY | 13
%01_DELAY | 14
%02_DELAY | 15
%03_DELAY | 14
%04_DELAY | 14
%05_DELAY | 14
%06_DELAY | 14
%07_DELAY | 13
%08_DELAY | 15
%09_DELAY | 15
%10_DELAY | 15
%11_DELAY | 16
#
%DEFECTS
#name | chip
#No defects found!
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | ResponseCurve
%TEST_DATE | 22/11/2001
%TEST_TIME | 10:03:50
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4660
%SCAN | 1
#
%HOST | ATLASPC9
%T0 | 34.0 | deg. C
%T1 | 35.0 | deg. C
#
#
%N_SCANS | 3
%Q_00 | 1.50 fC
%Q_01 | 2.00 fC
%Q_02 | 2.50 fC
%RC_FIT | 4 | (linear fit)
#
%ModuleResponseCurvePlot
# |  | func | p0 | p1 | p2
Chip 0: | 4 | 17.42 | 55.48 | 0.00
Chip 1: | 4 | 16.00 | 55.92 | 0.00
Chip 2: | 4 | 10.96 | 55.59 | 0.00
Chip 3: | 4 | 15.26 | 56.42 | 0.00
Chip 4: | 4 | 14.30 | 57.77 | 0.00
Chip 5: | 4 | 17.75 | 54.60 | 0.00
Chip 6: | 4 | 15.74 | 55.40 | 0.00
Chip 7: | 4 | 12.40 | 56.59 | 0.00
Chip 8: | 4 | 13.47 | 56.13 | 0.00
Chip 9: | 4 | 12.56 | 55.70 | 0.00
Chip 10: | 4 | 11.20 | 56.31 | 0.00
Chip 11: | 4 | 15.45 | 54.02 | 0.00
#
%ModuleGainOffsetNoise
# at 2.00fC
# |  | vt50 | rms | gain | rms | offset | rms | outnse | innse | rms | comment
Chip 0: | 127.9 | 7.48 | 55.5 | 0.93 | 18.2 | 6.87 | 5.68 | 640 | 20 | unbonded
Chip 1: | 127.6 | 7.48 | 55.9 | 1.00 | 16.3 | 7.74 | 5.67 | 634 | 20 | unbonded
Chip 2: | 122.5 | 6.71 | 55.6 | 1.14 | 12.0 | 5.63 | 5.28 | 594 | 20 | unbonded
Chip 3: | 128.2 | 7.42 | 56.4 | 1.02 | 15.8 | 6.65 | 5.67 | 628 | 19 | unbonded
Chip 4: | 130.0 | 7.79 | 57.8 | 1.01 | 15.0 | 6.80 | 5.52 | 597 | 19 | unbonded
Chip 5: | 126.7 | 8.16 | 54.6 | 0.93 | 18.1 | 7.79 | 5.56 | 636 | 20 | unbonded
Chip 6: | 126.4 | 7.55 | 55.4 | 0.92 | 16.2 | 7.44 | 5.61 | 633 | 20 | unbonded
Chip 7: | 125.6 | 6.57 | 56.6 | 0.98 | 13.1 | 6.48 | 5.68 | 627 | 19 | unbonded
Chip 8: | 125.4 | 7.71 | 56.1 | 0.87 | 14.0 | 7.05 | 5.62 | 626 | 20 | unbonded
Chip 9: | 124.2 | 8.38 | 55.7 | 0.89 | 14.2 | 7.28 | 5.55 | 622 | 19 | unbonded
Chip 10: | 123.7 | 7.90 | 56.3 | 1.04 | 12.0 | 7.27 | 5.57 | 619 | 17 | unbonded
Chip 11: | 123.0 | 7.44 | 54.0 | 1.01 | 15.7 | 7.30 | 5.55 | 643 | 25 | unbonded
#
#BadChannelSummary - not for the database
# at 2.00fC
# |  | lost | dodgy | dead | stuck | ineff | unbon | lo_gn | hi_gn | lo_off | hi_off | partbon | hi_nse
#Chip 0: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 1: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 2: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 3: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 4: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 5: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 6: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 7: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 8: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 9: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 10: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 11: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Link 0: | 768 | 0 | 0 | 0 | 0 | 768 | 0 | 0 | 0 | 0 | 0 | 0
#Link 1: | 768 | 0 | 0 | 0 | 0 | 768 | 0 | 0 | 0 | 0 | 0 | 0
#Link 2: | 1536 | 0 | 0 | 0 | 0 | 1536 | 0 | 0 | 0 | 0 | 0 | 0
#
%DEFECTS
#name | channel | value
#No defects found!
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 22/11/2001
%TEST_TIME | 10:18:51
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 4661
%SCAN | 1
#
%HOST | ATLASPC9
%T0 | 38.0 | deg. C
%T1 | 38.0 | deg. C
#
%TRIM_TYPE | 0
%TRIM_Q | 10.0
%N_SCANS | 16
%T_00 | 0
%T_01 | 1
%T_02 | 2
%T_03 | 3
%T_04 | 4
%T_05 | 5
%T_06 | 6
%T_07 | 7
%T_08 | 8
%T_09 | 9
%T_10 | 10
%T_11 | 11
%T_12 | 12
%T_13 | 13
%T_14 | 14
%T_15 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005hb220s30_tr0_20011122.trim
%MASK_FILE | D:\sctvar\results\0170200005hb220s30_tr0_20011122.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 0 | 92.5 | 128 | 92.2 | 1.22 | -18.3 | 2.0 | 3.87 | 0.17
Chip 1: | 0 | 92.5 | 128 | 92.2 | 1.18 | -19.8 | 2.6 | 3.54 | 0.17
Chip 2: | 0 | 92.5 | 128 | 91.7 | 0.97 | -22.4 | 2.7 | 3.00 | 0.26
Chip 3: | 0 | 92.5 | 128 | 91.9 | 1.25 | -18.6 | 2.0 | 3.79 | 0.16
Chip 4: | 0 | 92.5 | 128 | 91.7 | 1.17 | -20.4 | 2.1 | 3.51 | 0.19
Chip 5: | 0 | 92.5 | 128 | 92.2 | 1.08 | -20.8 | 2.6 | 3.37 | 0.11
Chip 6: | 0 | 92.5 | 128 | 91.9 | 1.13 | -19.5 | 2.2 | 3.56 | 0.16
Chip 7: | 0 | 92.5 | 128 | 91.7 | 1.09 | -19.7 | 2.1 | 3.51 | 0.17
Chip 8: | 0 | 92.5 | 128 | 92.0 | 1.03 | -19.9 | 2.4 | 3.44 | 0.15
Chip 9: | 0 | 92.5 | 128 | 92.0 | 1.08 | -20.3 | 2.5 | 3.33 | 0.14
Chip 10: | 0 | 92.5 | 128 | 92.0 | 1.09 | -19.3 | 2.3 | 3.45 | 0.14
Chip 11: | 0 | 92.5 | 127 | 92.1 | 1.30 | -16.7 | 1.8 | 4.11 | 0.36
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -11.34
%TR_STEP | 127 | 4.43
%TR_STEP | 165 | 4.64
%TR_OFFSET | 452 | -11.71
%TR_STEP | 511 | 4.40
%TR_OFFSET | 608 | -13.92
%TR_STEP | 639 | 4.18
%TR_OFFSET | 210 | -26.19
%TR_STEP | 249 | 4.03
%TR_STEP | 251 | 4.10
%TR_STEP | 255 | 4.13
%TR_OFFSET | 585 | -26.48
%TR_OFFSET | 750 | -25.16
%TR_STEP | 765 | 5.22
#14 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 22/11/2001
%TEST_TIME | 10:18:51
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 4661
%SCAN | 17
#
%HOST | ATLASPC9
%T0 | 38.0 | deg. C
%T1 | 38.0 | deg. C
#
%TRIM_TYPE | 1
%TRIM_Q | 10.0
%N_SCANS | 4
%T_00 | 3
%T_01 | 7
%T_02 | 11
%T_03 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005hb220s30_tr1_20011122.trim
%MASK_FILE | D:\sctvar\results\0170200005hb220s30_tr1_20011122.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 1 | 115.0 | 128 | 115.2 | 2.18 | -10.0 | 1.0 | 7.17 | 0.30
Chip 1: | 1 | 115.0 | 128 | 115.3 | 1.94 | -10.5 | 1.4 | 6.69 | 0.42
Chip 2: | 1 | 115.0 | 128 | 114.8 | 1.77 | -10.8 | 1.3 | 6.14 | 0.43
Chip 3: | 1 | 115.0 | 128 | 115.0 | 2.04 | -9.9 | 1.1 | 7.15 | 0.27
Chip 4: | 1 | 115.0 | 128 | 114.7 | 1.90 | -10.6 | 1.2 | 6.78 | 0.31
Chip 5: | 1 | 115.0 | 128 | 115.2 | 1.84 | -11.0 | 1.3 | 6.43 | 0.22
Chip 6: | 1 | 115.0 | 128 | 115.2 | 1.97 | -10.3 | 1.2 | 6.78 | 0.27
Chip 7: | 1 | 115.0 | 128 | 115.4 | 1.98 | -9.9 | 1.1 | 6.95 | 0.30
Chip 8: | 1 | 115.0 | 128 | 114.7 | 2.06 | -10.2 | 1.2 | 6.67 | 0.27
Chip 9: | 1 | 115.0 | 128 | 115.2 | 1.86 | -10.4 | 1.4 | 6.49 | 0.24
Chip 10: | 1 | 115.0 | 128 | 115.2 | 2.00 | -9.9 | 1.2 | 6.72 | 0.25
Chip 11: | 1 | 115.0 | 128 | 114.7 | 2.26 | -9.0 | 0.9 | 7.70 | 0.57
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -6.13
%TR_STEP | 165 | 10.58
%TR_OFFSET | 452 | -5.78
%TR_STEP | 511 | 8.06
%TR_STEP | 639 | 7.94
%TR_OFFSET | 691 | -6.75
%TR_OFFSET | 734 | -6.71
%TR_OFFSET | 171 | -13.43
%TR_OFFSET | 210 | -13.41
%TR_STEP | 251 | 7.88
%TR_STEP | 255 | 7.94
%TR_OFFSET | 585 | -13.67
%TR_OFFSET | 750 | -12.44
%TR_STEP | 765 | 9.54
#14 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 22/11/2001
%TEST_TIME | 10:18:51
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 4661
%SCAN | 21
#
%HOST | ATLASPC9
%T0 | 38.0 | deg. C
%T1 | 38.0 | deg. C
#
%TRIM_TYPE | 2
%TRIM_Q | 10.0
%N_SCANS | 4
%T_00 | 3
%T_01 | 7
%T_02 | 11
%T_03 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005hb220s30_tr2_20011122.trim
%MASK_FILE | D:\sctvar\results\0170200005hb220s30_tr2_20011122.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 2 | 137.5 | 128 | 136.7 | 3.28 | -6.7 | 0.6 | 10.64 | 0.46
Chip 1: | 2 | 137.5 | 128 | 137.2 | 2.52 | -7.1 | 1.0 | 9.94 | 0.62
Chip 2: | 2 | 137.5 | 128 | 137.0 | 2.74 | -7.1 | 0.9 | 9.27 | 0.57
Chip 3: | 2 | 137.5 | 128 | 136.9 | 3.14 | -6.6 | 0.7 | 10.67 | 0.40
Chip 4: | 2 | 137.5 | 128 | 137.4 | 2.84 | -7.1 | 0.8 | 10.10 | 0.44
Chip 5: | 2 | 137.5 | 128 | 136.4 | 2.79 | -7.4 | 0.8 | 9.59 | 0.33
Chip 6: | 2 | 137.5 | 128 | 136.5 | 3.00 | -6.9 | 0.8 | 10.10 | 0.41
Chip 7: | 2 | 137.5 | 128 | 136.8 | 3.14 | -6.5 | 0.8 | 10.44 | 0.42
Chip 8: | 2 | 137.5 | 128 | 137.0 | 2.86 | -6.9 | 0.8 | 9.89 | 0.39
Chip 9: | 2 | 137.5 | 128 | 136.9 | 2.78 | -7.0 | 0.9 | 9.65 | 0.37
Chip 10: | 2 | 137.5 | 128 | 136.9 | 2.95 | -6.6 | 0.8 | 9.95 | 0.38
Chip 11: | 2 | 137.5 | 128 | 136.4 | 3.28 | -6.2 | 0.7 | 11.16 | 0.78
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -4.27
%TR_STEP | 165 | 15.63
%TR_OFFSET | 452 | -4.01
%TR_STEP | 511 | 11.99
%TR_STEP | 639 | 11.54
%TR_OFFSET | 691 | -4.64
%TR_OFFSET | 734 | -4.63
%TR_OFFSET | 171 | -8.89
%TR_OFFSET | 210 | -9.04
%TR_OFFSET | 585 | -9.26
%TR_OFFSET | 750 | -9.24
%TR_STEP | 765 | 13.65
#12 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 22/11/2001
%TEST_TIME | 10:18:51
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 4661
%SCAN | 25
#
%HOST | ATLASPC9
%T0 | 38.0 | deg. C
%T1 | 38.0 | deg. C
#
%TRIM_TYPE | 3
%TRIM_Q | 10.0
%N_SCANS | 4
%T_00 | 3
%T_01 | 7
%T_02 | 11
%T_03 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005hb220s30_tr3_20011122.trim
%MASK_FILE | D:\sctvar\results\0170200005hb220s30_tr3_20011122.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 3 | 162.5 | 128 | 162.6 | 3.88 | -5.0 | 0.5 | 14.09 | 0.55
Chip 1: | 3 | 162.5 | 128 | 161.1 | 3.94 | -5.2 | 0.8 | 13.29 | 0.80
Chip 2: | 3 | 162.5 | 128 | 162.2 | 3.71 | -5.4 | 0.6 | 12.36 | 0.71
Chip 3: | 3 | 162.5 | 128 | 162.0 | 4.11 | -4.9 | 0.5 | 14.11 | 0.51
Chip 4: | 3 | 162.5 | 128 | 161.8 | 3.90 | -5.4 | 0.6 | 13.31 | 0.56
Chip 5: | 3 | 162.5 | 128 | 162.6 | 3.81 | -5.4 | 0.6 | 12.91 | 0.44
Chip 6: | 3 | 162.5 | 128 | 162.6 | 3.92 | -5.1 | 0.5 | 13.46 | 0.53
Chip 7: | 3 | 162.5 | 128 | 161.8 | 4.04 | -5.0 | 0.6 | 13.78 | 0.55
Chip 8: | 3 | 162.5 | 128 | 162.7 | 3.79 | -5.1 | 0.6 | 13.16 | 0.50
Chip 9: | 3 | 162.5 | 128 | 162.0 | 3.77 | -5.2 | 0.6 | 12.84 | 0.48
Chip 10: | 3 | 162.5 | 128 | 161.9 | 3.88 | -5.0 | 0.6 | 13.23 | 0.50
Chip 11: | 3 | 162.5 | 128 | 162.4 | 4.31 | -4.6 | 0.5 | 14.76 | 0.91
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -3.26
%TR_STEP | 165 | 20.78
%TR_STEP | 258 | 14.59
%TR_OFFSET | 452 | -3.13
%TR_STEP | 511 | 15.65
%TR_OFFSET | 608 | -3.51
%TR_STEP | 639 | 15.13
%TR_OFFSET | 734 | -3.43
%TR_OFFSET | 171 | -6.78
%TR_OFFSET | 210 | -6.91
%TR_OFFSET | 585 | -6.89
%TR_STEP | 627 | 14.72
%TR_OFFSET | 750 | -8.11
%TR_STEP | 765 | 17.64
#14 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 22/11/2001
%TEST_TIME | 10:18:51
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 4661
%SCAN | 1
#
%HOST | ATLASPC9
%T0 | 38.0 | deg. C
%T1 | 38.0 | deg. C
#
%TRIM_TYPE | -1
%TRIM_Q | 10.0
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005hb220s30_tr-1_20011122.trim
%MASK_FILE | D:\sctvar\results\0170200005hb220s30_tr-1_20011122.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 0 | 95.0 | 128 | 95.0 | 1.09 | -18.3 | 2.0 | 3.87 | 0.17
Chip 1: | 0 | 95.0 | 128 | 95.2 | 1.04 | -19.8 | 2.6 | 3.54 | 0.17
Chip 2: | 1 | 95.0 | 128 | 95.2 | 1.79 | -10.8 | 1.3 | 6.14 | 0.43
Chip 3: | 0 | 95.0 | 128 | 94.8 | 1.04 | -18.6 | 2.0 | 3.79 | 0.16
Chip 4: | 0 | 95.0 | 128 | 95.1 | 0.94 | -20.4 | 2.1 | 3.51 | 0.19
Chip 5: | 0 | 95.0 | 128 | 95.0 | 0.97 | -20.8 | 2.6 | 3.37 | 0.11
Chip 6: | 0 | 95.0 | 128 | 94.9 | 1.03 | -19.5 | 2.2 | 3.56 | 0.16
Chip 7: | 0 | 95.0 | 128 | 95.1 | 1.02 | -19.7 | 2.1 | 3.51 | 0.17
Chip 8: | 0 | 95.0 | 128 | 95.2 | 0.91 | -19.9 | 2.4 | 3.44 | 0.15
Chip 9: | 0 | 95.0 | 128 | 95.0 | 0.98 | -20.3 | 2.5 | 3.33 | 0.14
Chip 10: | 0 | 95.0 | 128 | 95.0 | 1.05 | -19.3 | 2.3 | 3.45 | 0.14
Chip 11: | 1 | 95.0 | 128 | 95.1 | 2.43 | -9.0 | 0.9 | 7.70 | 0.57
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -11.34
%TR_STEP | 127 | 4.43
%TR_STEP | 165 | 4.64
%TR_OFFSET | 452 | -11.71
%TR_STEP | 511 | 4.40
%TR_OFFSET | 608 | -13.92
%TR_STEP | 639 | 4.18
%TR_OFFSET | 210 | -26.19
%TR_STEP | 249 | 4.03
%TR_STEP | 251 | 4.10
%TR_STEP | 255 | 4.13
%TR_OFFSET | 585 | -26.48
%TR_OFFSET | 750 | -12.44
%TR_STEP | 765 | 9.54
#14 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30q1t092r0LT
%TEST_NAME | ResponseCurve
%TEST_DATE | 22/11/2001
%TEST_TIME | 10:50:01
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4664
%SCAN | 1
#
%HOST | ATLASPC9
%T0 | 34.0 | deg. C
%T1 | 34.0 | deg. C
#
#
%N_SCANS | 10
%Q_00 | 0.50 fC
%Q_01 | 0.75 fC
%Q_02 | 1.00 fC
%Q_03 | 1.25 fC
%Q_04 | 1.50 fC
%Q_05 | 2.00 fC
%Q_06 | 3.00 fC
%Q_07 | 4.00 fC
%Q_08 | 6.00 fC
%Q_09 | 8.00 fC
%RC_FIT | 3 | (exponential fit)
#
%ModuleResponseCurvePlot
# |  | func | p0 | p1 | p2
Chip 0: | 3 | 1457.52 | 6.25 | -694.93
Chip 1: | 3 | 1412.38 | 6.10 | -671.65
Chip 2: | 3 | 1251.80 | 5.55 | -590.08
Chip 3: | 3 | 1417.84 | 6.10 | -674.74
Chip 4: | 3 | 1428.09 | 6.08 | -680.24
Chip 5: | 3 | 1348.48 | 5.92 | -638.86
Chip 6: | 3 | 1388.71 | 6.08 | -659.05
Chip 7: | 3 | 1451.00 | 6.38 | -689.76
Chip 8: | 3 | 1355.38 | 5.87 | -642.85
Chip 9: | 3 | 1362.76 | 6.02 | -645.35
Chip 10: | 3 | 1338.89 | 5.91 | -632.93
Chip 11: | 3 | 1610.31 | 7.26 | -768.34
#
%ModuleGainOffsetNoise
# at 2.00fC
# |  | vt50 | rms | gain | rms | offset | rms | outnse | innse | rms | comment
Chip 0: | 148.7 | 1.52 | 56.9 | 0.98 | 33.8 | 1.73 | 5.46 | 601 | 21 | unbonded
Chip 1: | 149.1 | 1.49 | 56.4 | 1.00 | 34.4 | 1.70 | 5.41 | 599 | 18 | unbonded
Chip 2: | 148.2 | 1.37 | 54.6 | 1.10 | 36.0 | 1.56 | 5.21 | 597 | 19 | unbonded
Chip 3: | 149.3 | 1.64 | 56.5 | 1.13 | 34.2 | 1.72 | 5.40 | 598 | 20 | unbonded
Chip 4: | 150.5 | 1.50 | 57.1 | 1.13 | 33.9 | 1.68 | 5.27 | 577 | 18 | unbonded
Chip 5: | 147.9 | 1.32 | 55.4 | 1.07 | 35.4 | 1.74 | 5.25 | 593 | 18 | unbonded
Chip 6: | 148.4 | 1.34 | 55.6 | 0.93 | 35.3 | 1.61 | 5.32 | 598 | 19 | unbonded
Chip 7: | 149.0 | 1.41 | 55.5 | 0.92 | 35.9 | 1.42 | 5.48 | 618 | 17 | unbonded
Chip 8: | 149.1 | 1.34 | 56.0 | 1.00 | 34.9 | 1.45 | 5.45 | 608 | 17 | unbonded
Chip 9: | 148.6 | 1.15 | 55.2 | 0.91 | 35.9 | 1.67 | 5.28 | 598 | 20 | unbonded
Chip 10: | 149.2 | 1.36 | 55.1 | 1.06 | 36.5 | 1.66 | 5.39 | 611 | 19 | unbonded
Chip 11: | 146.7 | 1.63 | 54.4 | 1.23 | 36.8 | 1.90 | 5.40 | 620 | 19 | unbonded
#
#BadChannelSummary - not for the database
# at 2.00fC
# |  | lost | dodgy | dead | stuck | ineff | unbon | lo_gn | hi_gn | lo_off | hi_off | partbon | hi_nse
#Chip 0: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 1: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 2: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 3: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 4: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 5: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 6: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 7: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 8: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 9: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 10: | 128 | 0 | 0 | 0 | 0 | 128 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 11: | 128 | 0 | 1 | 0 | 0 | 127 | 0 | 0 | 0 | 0 | 0 | 0
#Link 0: | 768 | 0 | 0 | 0 | 0 | 768 | 0 | 0 | 0 | 0 | 0 | 0
#Link 1: | 768 | 0 | 1 | 0 | 0 | 767 | 0 | 0 | 0 | 0 | 0 | 0
#Link 2: | 1536 | 0 | 1 | 0 | 0 | 1535 | 0 | 0 | 0 | 0 | 0 | 0
#
%DEFECTS
#name | channel | value
%DEAD | 1518
#1 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005hb220s30q1t092r0LT
%TEST_NAME | Timewalk
%TEST_DATE | 22/11/2001
%TEST_TIME | 11:04:30
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4665
%SCAN | 2
#
%HOST | ATLASPC9
%T0 | 36.0 | deg. C
%T1 | 38.0 | deg. C
#
#
# run scan charge
# 4665 2 1.25
# 4665 3 1.50
# 4665 4 1.75
# 4665 5 2.00
# 4665 6 3.00
# 4665 7 4.00
# 4665 8 5.00
# 4665 9 6.00
# 4665 10 8.00
# 4665 11 10.00
#
%00_TW | 9.6
%01_TW | 9.9
%02_TW | 10.1
%03_TW | 10.2
%04_TW | 10.1
%05_TW | 10.0
%06_TW | 10.1
%07_TW | 10.4
%08_TW | 10.2
%09_TW | 10.2
%10_TW | 10.1
%11_TW | 9.2
#
%00_TCAL | 28.6
%01_TCAL | 28.4
%02_TCAL | 29.9
%03_TCAL | 28.8
%04_TCAL | 28.6
%05_TCAL | 28.6
%06_TCAL | 28.6
%07_TCAL | 28.3
%08_TCAL | 28.9
%09_TCAL | 28.3
%10_TCAL | 28.1
%11_TCAL | 28.7
#
%DEFECTS
#name | chip | timewalk
#No defects found!

## MODULE
atlasPartsId | 20220170200005
manufacturerSerialNo | <data>
eventDescription | MODULE
dateEntry [dd/mm/yyyy] | 08/12/2001
location [instituteCode(DB)] | KEK
personInitial
problem [YES/NO] | NO
pass [YES/NO] | YES
comments
SENSOR_BASEBOARD | 20220080100096
ASIC_HYBRID | 20220170200005
dateAssemblyHyb1 [dd/mm/yyyy] | 22/11/2001
location [instituteCode(DB)] | KEK
personInitial
assemblyJigID | SB1,H2
ARALDITE2011
BN FILLER
assemblyTemperature(xx.x) [C] | 25
cureTemperature(xx) [C] | 25
cureDurationTime [hrs] | 15
dateAssemblyHyb2 [dd/mm/yyyy] | 23/11/2001
location [instituteCode(DB)] | HAM
personInitial
assemblyJigID | SB1,H2
ARALDITE2011
BN FILLER
assemblyTemperature(xx.x) [C] | 25
cureTemperature(xx) [C] | 25
cureDurationTime [hrs] | 24
dateWireBonding [dd/mm/yyyy] | 26/11/2001
location [instituteCode(DB)] | HAM
nDefectiveChannels
defectiveChannels
comment
dateSURVEY_XY-INITIAL | 27/11/2001
dateSURVEY_Z-INITIAL | 27/11/2001
dateI_V-INITIAL | 27/11/2001
dateELECTRICAL-INITIAL | 27/11/2001
dateTHERMALCYCLE | 29/11/2001
loTemperature [C] | -30
hiTemperature [C] | +50
numberRepetition | 10
dateSURVEY_XY-TC | 30/11/2001
dateSURVEY_Z-TC | 30/11/2001
dateI_V-TC | 30/11/2001
dateELECTRICAL-TC | 30/11/2001
dateLONGTERM_MODULE | 30/11/2001
dateI_STABILITY-LT | 02/12/2001
dateSURVEY_XY-LT | 03/12/2001
dateSURVEY_Z-LT | 03/12/2001
dateI_V-LT | 06/12/2001
dateELECTRICAL-LT | 03/12/2001
dateIRRAD
dateANNEAL
dateLASER
dateSOURCE

## I_V-INITIAL
atlasPartsId | 20220170200005
eventDescription | I_V-INITIAL | X: SB, INITIAL, TC, LT, ANNEAL, ...
dateEntry [dd/mm/yyyy] | 27/11/2001
location [instituteCode(DB)] | KEK
personInitial | ta
problem [YES/NO] | no
pass [YES/NO] | yes
comments
temperature(xx.x) [C] | 15.41
moduleI_150V [microA] | 0.1319
moduleI_350V [microA] | 0.1832
moduleI_500V [microA] | 0.2147
temperature(xx.x) [C] | <data>
LeftUpperI_150V [microA] | <data> | : optional
LeftUpperI_350V [microA] | <data>
LeftUpperI_500V [microA] | <data>
temperature(xx.x) [C] | <data>
RightUpperI_150V [microA] | <data> | : optional
RightUpperI_350V [microA] | <data>
RightUpperI_500V [microA] | <data>
temperature(xx.x) [C] | <data>
LeftLowerI_150V [microA] | <data> | : optional
LeftLowerI_350V [microA] | <data>
LeftLowerI_500V [microA] | <data>
temperature(xx.x) [C] | <data>
RightLowerI_150V [microA] | <data> | : optional
RightLowerI_350V [microA] | <data>
RightLowerI_500V [microA] | <data>
RAW DATA
Bias [V] | I_total [microA]
0.000000 | -0.00187
10.000000 | 0.04877
20.000000 | 0.06678
30.000000 | 0.07798
40.000000 | 0.08649
50.000000 | 0.09328
60.000000 | 0.09904
70.000000 | 0.10353
80.000000 | 0.10786
90.000000 | 0.11176
100.000000 | 0.11572
110.000000 | 0.11893
120.000000 | 0.12235
130.000000 | 0.12633
140.000000 | 0.12869
150.000000 | 0.13192
160.000000 | 0.13444
170.000000 | 0.13767
180.000000 | 0.14001
190.000000 | 0.1433
200.000000 | 0.14727
210.000000 | 0.15035
220.000000 | 0.15198
230.000000 | 0.15486
240.000000 | 0.15756
250.000000 | 0.15976
260.000000 | 0.16269
270.000000 | 0.1651
280.000000 | 0.1679
290.000000 | 0.16918
300.000000 | 0.17164
310.000000 | 0.17442
320.000000 | 0.17702
330.000000 | 0.1789
340.000000 | 0.18129
350.000000 | 0.18315
360.000000 | 0.18555
370.000000 | 0.18768
380.000000 | 0.18916
390.000000 | 0.19215
400.000000 | 0.19389
410.000000 | 0.19622
420.000000 | 0.19818
430.000000 | 0.20017
440.000000 | 0.20236
450.000000 | 0.20431
460.000000 | 0.20595
470.000000 | 0.2088
480.000000 | 0.21041
490.000000 | 0.2122
500.000000 | 0.21474

## SURVEY_XY-INITIAL
atlasPartsId | 20220170200005
eventDescription | SURVEY_XY-INITIAL | : X=SB, INITIAL, TC, LT, ...
date [dd/mm/yyyy] | 17/7/2001
location [instituteCode(DB)] | KEK
personInitials | TKr
problem [YES/NO] | NO
pass [YES/NO] | YES
comments | any comments
temperature(xx.x) [C] | 25.6
measurementJigID | <data>
mhx [um] | 6.9184
mhy [um] | -11.0213
msx [um] | 4.8538
msy [um] | -7.2639
sepf [um] | 2.9547
sepb [um] | -0.55045
midxf [um] | 0.49064
midyf [um] | 2.8611
a1 [mrad] | 0.0090401
a2 [mrad] | -0.081992
a3 [mrad] | 0.0021814
a4 [mrad] | 0.0008119
stereo [mrad] | 0.02953
hymxf [um] | 67.0506
hymyf [um] | -8.8237
hymaf [mrad] | 0.50826
hymxb [um] | -4.357
hymyb [um] | 8.1685
hymab [mrad] | -0.86623
conp1x [um] | 102.3497
conp1y [um] | -6.979

## SURVEY_Z-INITIAL
atlasPartsId | 20220170200005
eventDescription | SURVEY_Z-INITIAL | : X=INITIAL, TC, LT, ...
date [dd/mm/yyy] | 6/12/2001
location [instituteCode(DB)] | KEK
personInitial | TKr
problem [YES/NO] | NO
pass [YES/NO] | YES
comments | V2.2.0
temperature [C] | 26.5
measurementJigID | <data>
maxZlower [mm] | -0.033403
maxZupper [mm] | 0.0093752
midplaneEq | z=ax+by+c
Left a | 0.00040215
b | 0.00044151
c | 0.47357
Right a | -0.00078795
b | 0.000236
c | 0.4705
midplaneHeight [mm] | 0.47203
moduleThickness [mm] | 1.1797
optimalMaxZerrorLower [mm] | 0.032249
optimalMaxZerrorUpper [mm] | 0.033577
optimalRmsZerrorLower [mm] | 0.0088962
optimalRmsZerrorUpper [mm] | 0.0083856
moduleConcavity x [mm] | 0.03789
y | -0.00088343
sensorSkew x [mm] | -0.025524
y | 0.021705
coolingTabThickness [mm] | 0.93443
farTabThickness [mm] | 0.95865
halfTabThickness [mm] | 0.47327
tabSkew y [mm] | -0.012109
adhesiveThicknessTotal [mm] | 0.16313
adhesiveAsymmetry [mm] | 0.0012386
loCoolingFacing a [mrad] | -0.016576
b [mrad] | -1.5135
loCoolingFacingConcavity [mm] | 0
optimalFrameRawDATA
leftLower | x
1 | -63.1983
2 | -47.3269
3 | -31.4552
4 | -15.5838
5 | 0.28779
6 | -63.5156
7 | -47.6442
8 | -31.7726
9 | -15.9011
10 | -0.029681
11 | -63.833
12 | -47.9616
13 | -32.0901
14 | -16.2186
15 | -0.34706
16 | -64.1506
17 | -48.279
18 | -32.4076
19 | -16.536
20 | -0.66444
21 | -64.4679
22 | -48.5965
23 | -32.725
24 | -16.8534
25 | -0.98193
rightLower | x
1 | 0.98173
2 | 16.8539
3 | 32.7247
4 | 48.5964
5 | 64.4679
6 | 0.66426
7 | 16.5368
8 | 32.4074
9 | 48.2789
10 | 64.1504
11 | 0.34687
12 | 16.2191
13 | 32.09
14 | 47.9615
15 | 63.8332
16 | 0.029597
17 | 15.9018
18 | 31.7725
19 | 47.644
20 | 63.5156
21 | -0.28799
22 | 15.5841
23 | 31.4551
24 | 47.3266
25 | 63.1981
leftUpper | x
1 | -63.1513
2 | -47.2798
3 | -31.4081
4 | -15.5364
5 | 0.33533
6 | -63.467
7 | -47.5953
8 | -31.7236
9 | -15.8521
10 | 0.01965
11 | -63.7826
12 | -47.9109
13 | -32.0394
14 | -16.1677
15 | -0.29604
16 | -64.0982
17 | -48.2266
18 | -32.355
19 | -16.4833
20 | -0.61173
21 | -64.414
22 | -48.5424
23 | -32.6707
24 | -16.799
25 | -0.92722
rightUpper | x
1 | 0.92728
2 | 16.799
3 | 32.6706
4 | 48.5423
5 | 64.414
6 | 0.6116
7 | 16.4834
8 | 32.3551
9 | 48.2266
10 | 64.0983
11 | 0.29601
12 | 16.1677
13 | 32.0393
14 | 47.9109
15 | 63.7827
16 | -0.019577
17 | 15.8523
18 | 31.7237
19 | 47.5952
20 | 63.467
21 | -0.33537
22 | 15.5364
23 | 31.408
24 | 47.2798
25 | 63.1512
leftLower | y
1 | -32.8276
2 | -32.5079
3 | -32.1883
4 | -31.8685
5 | -31.5488
6 | -17.0565
7 | -16.7365
8 | -16.4168
9 | -16.097
10 | -15.7774
11 | -1.2849
12 | -0.96505
13 | -0.64533
14 | -0.3256
15 | -0.0058423
16 | 14.4867
17 | 14.8064
18 | 15.1261
19 | 15.4458
20 | 15.7656
21 | 30.2581
22 | 30.5778
23 | 30.8975
24 | 31.2172
25 | 31.5371
rightLower | y
1 | -31.5365
2 | -31.2188
3 | -30.8971
4 | -30.5773
5 | -30.2575
6 | -15.7651
7 | -15.4484
8 | -15.1257
9 | -14.8059
10 | -14.4861
11 | 0.006432
12 | 0.32402
13 | 0.64572
14 | 0.96554
15 | 1.2853
16 | 15.7778
17 | 16.0946
18 | 16.4173
19 | 16.737
20 | 17.0569
21 | 31.5492
22 | 31.867
23 | 32.1886
24 | 32.5084
25 | 32.8283
leftUpper | y
1 | -32.8214
2 | -32.5037
3 | -32.186
4 | -31.8682
5 | -31.5506
6 | -17.0506
7 | -16.7328
8 | -16.4152
9 | -16.0975
10 | -15.7798
11 | -1.2798
12 | -0.96195
13 | -0.64428
14 | -0.32663
15 | -0.008886
16 | 14.491
17 | 14.8089
18 | 15.1265
19 | 15.4442
20 | 15.7619
21 | 30.2619
22 | 30.5797
23 | 30.8974
24 | 31.2149
25 | 31.5328
rightUpper | y
1 | -31.533
2 | -31.2137
3 | -30.8977
4 | -30.5801
5 | -30.2626
6 | -15.7622
7 | -15.4419
8 | -15.1269
9 | -14.8093
10 | -14.4917
11 | 0.0086818
12 | 0.32805
13 | 0.64392
14 | 0.96162
15 | 1.2792
16 | 15.7795
17 | 16.0996
18 | 16.4148
19 | 16.7324
20 | 17.05
21 | 31.5504
22 | 31.8697
23 | 32.1855
24 | 32.5031
25 | 32.8207
leftLower | z
1 | 0.084596
2 | 0.039124
3 | 0.019266
4 | 0.013061
5 | 0.015172
6 | 0.058998
7 | 0.016227
8 | 0.0037135
9 | 0.00085657
10 | 0.0034916
11 | 0.052198
12 | 0.011765
13 | 0.00067378
14 | -0.0019982
15 | 0.0067409
16 | 0.060434
17 | 0.015752
18 | 0.00041005
19 | -0.004512
20 | -0.0056369
21 | 0.087377
22 | 0.03747
23 | 0.013856
24 | 0.0052523
25 | 0.0036584
rightLower | z
1 | 0.0098568
2 | 0.01341
3 | 0.016964
4 | 0.039185
5 | 0.084826
6 | 0.00087045
7 | -3.7437e-05
8 | -0.00094445
9 | 0.013716
10 | 0.05588
11 | 0.0047032
12 | 0.00020749
13 | -0.0042877
14 | 0.0056939
15 | 0.044518
16 | 0.00063705
17 | -0.0014615
18 | -0.003559
19 | 0.010561
20 | 0.050933
21 | 0.012112
22 | 0.011731
23 | 0.011352
24 | 0.032921
25 | 0.079183
leftUpper | z
1 | -0.093376
2 | -0.044965
3 | -0.019872
4 | -0.010128
5 | -0.010828
6 | -0.06321
7 | -0.016082
8 | -0.0017025
9 | 0.0025322
10 | 0.0014778
11 | -0.051215
12 | -0.010828
13 | -0.00061403
14 | 0.00062356
15 | -0.0028028
16 | -0.056956
17 | -0.012102
18 | 0.0018463
19 | 0.0034429
20 | 0.0011296
21 | -0.082502
22 | -0.03443
23 | -0.013321
24 | -0.0088068
25 | -0.015255
rightUpper | z
1 | -0.012685
2 | -0.014774
3 | -0.016864
4 | -0.033646
5 | -0.073504
6 | -6.3639e-05
7 | -0.0015956
8 | -0.0031298
9 | -0.015408
10 | -0.053528
11 | -0.0048407
12 | -0.0026892
13 | -0.00053911
14 | -0.010851
15 | -0.048398
16 | 0.0041501
17 | 0.0020831
18 | 1.38e-05
19 | -0.013798
20 | -0.054929
21 | -0.0066751
22 | -0.0085915
23 | -0.010509
24 | -0.032456
25 | -0.075734

## ELECTRICAL-INITIAL
atlasPartsId | 20220170200005
eventDescription | ELECTRICAL-INITIAL |  | : X=INITIAL, TC, LT, ANNEAL, ...
date [dd/mm/yyyy] | 27/11/2001
location [instituteCode(DB)] | KEK
personInitial | ta
problem [YES/NO] | no
pass [YES/NO] | yes
comments
#
# | http://hepwww.rl.ac.uk/atlas-sct/documents/dev/Electrical_Tests_301.doc
# | N.B. Some revisions have been incorporated here, but not yet in the reference,
# | namely that CONFIGURATION will be stored as an ASCII file and
# | and that the RUN field will be used to store both RUN and SCAN number.
#
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30
%TEST_NAME | FullBypassTest
%TEST_DATE | 27/11/2001
%TEST_TIME | 19:47:22
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4785
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 28.0 | deg. C
%T1 | 27.0 | deg. C
#
#NConfigs 37
#Vdd = 4.2 - 3.3
#
%M0->S1: | 3.3 V | (Minimal tested)
%M0->S2: | 3.3 V | (Minimal tested)
%S1->S2: | 3.3 V | (Minimal tested)
%S1->S3: | 3.3 V | (Minimal tested)
%S2->S3: | 3.3 V | (Minimal tested)
%S2->S4: | 3.3 V | (Minimal tested)
%S3->S4: | 3.3 V | (Minimal tested)
%S3->E5: | 3.3 V | (Minimal tested)
%S4->E5: | 3.3 V | (Minimal tested)
%E5->S9: | 3.3 V | (Minimal tested)
%M8->S9: | 3.3 V | (Minimal tested)
%M8->S10: | 3.3 V | (Minimal tested)
%S9->S10: | 3.3 V | (Minimal tested)
%S9->S11: | 3.3 V | (Minimal tested)
%S10->S11: | 3.3 V | (Minimal tested)
%S10->S12: | 3.3 V | (Minimal tested)
%S11->S12: | 3.3 V | (Minimal tested)
%S11->E13: | 3.3 V | (Minimal tested)
%S12->E13: | 3.3 V | (Minimal tested)
%E13->S1: | 3.3 V | (Minimal tested)
#
%DEFECTS
#name | bypass_link | Vdd_min
#No defects found!
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30
%TEST_NAME | PipelineTest
%TEST_DATE | 27/11/2001
%TEST_TIME | 19:47:47
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4785
%SCAN | 11
#
%HOST | ATLASPC11
%T0 | 28.0 | deg. C
%T1 | 28.0 | deg. C
#
#
%00_GOOD | 128
%01_GOOD | 128
%02_GOOD | 128
%03_GOOD | 128
%04_GOOD | 128
%05_GOOD | 128
%06_GOOD | 128
%07_GOOD | 128
%08_GOOD | 128
%09_GOOD | 128
%10_GOOD | 128
%11_GOOD | 128
#
%DEFECTS
#name | chan | cell
#No defects found!
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30
%TEST_NAME | StrobeDelay
%TEST_DATE | 27/11/2001
%TEST_TIME | 19:55:25
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4786
%SCAN | 2
#
%HOST | ATLASPC11
%T0 | 28.0 | deg. C
%T1 | 27.0 | deg. C
#
#
#parameter | value
%00_DELAY | 11
%01_DELAY | 12
%02_DELAY | 12
%03_DELAY | 11
%04_DELAY | 12
%05_DELAY | 12
%06_DELAY | 12
%07_DELAY | 11
%08_DELAY | 12
%09_DELAY | 12
%10_DELAY | 12
%11_DELAY | 14
#
%DEFECTS
#name | chip
#No defects found!
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30
%TEST_NAME | ResponseCurve
%TEST_DATE | 27/11/2001
%TEST_TIME | 20:07:31
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4787
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 27.0 | deg. C
%T1 | 27.0 | deg. C
#
#
%N_SCANS | 3
%Q_00 | 1.50 fC
%Q_01 | 2.00 fC
%Q_02 | 2.50 fC
%RC_FIT | 4 | (linear fit)
#
%ModuleResponseCurvePlot
# |  | func | p0 | p1 | p2
Chip 0: | 4 | 18.31 | 55.54 | 0.00
Chip 1: | 4 | 16.64 | 55.55 | 0.00
Chip 2: | 4 | 12.97 | 54.32 | 0.00
Chip 3: | 4 | 16.72 | 55.79 | 0.00
Chip 4: | 4 | 16.08 | 57.54 | 0.00
Chip 5: | 4 | 19.49 | 54.60 | 0.00
Chip 6: | 4 | 17.97 | 55.29 | 0.00
Chip 7: | 4 | 14.88 | 56.51 | 0.00
Chip 8: | 4 | 15.38 | 55.09 | 0.00
Chip 9: | 4 | 14.37 | 54.78 | 0.00
Chip 10: | 4 | 12.58 | 55.41 | 0.00
Chip 11: | 4 | 16.93 | 54.09 | 0.00
#
%ModuleGainOffsetNoise
# at 2.00fC
# |  | vt50 | rms | gain | rms | offset | rms | outnse | innse | rms | comment
Chip 0: | 129.1 | 7.54 | 55.5 | 0.99 | 18.8 | 6.97 | 13.40 | 1508 | 32 | OK
Chip 1: | 127.8 | 7.57 | 55.6 | 1.02 | 16.8 | 7.67 | 13.36 | 1503 | 27 | OK
Chip 2: | 122.6 | 8.19 | 54.2 | 3.76 | 13.6 | 5.81 | 13.69 | 1494 | 119 | OK
Chip 3: | 128.5 | 7.56 | 55.8 | 1.16 | 17.2 | 6.62 | 13.63 | 1527 | 35 | OK
Chip 4: | 131.4 | 8.02 | 57.5 | 1.19 | 16.5 | 6.79 | 13.64 | 1482 | 33 | OK
Chip 5: | 128.6 | 8.32 | 54.6 | 0.98 | 19.9 | 7.74 | 13.12 | 1502 | 39 | OK
Chip 6: | 128.6 | 7.54 | 55.3 | 1.06 | 18.1 | 7.31 | 13.22 | 1494 | 32 | OK
Chip 7: | 128.0 | 6.58 | 56.5 | 1.04 | 14.9 | 6.72 | 13.81 | 1528 | 37 | OK
Chip 8: | 125.5 | 7.88 | 55.1 | 0.99 | 15.8 | 6.95 | 13.38 | 1519 | 31 | OK
Chip 9: | 124.2 | 8.48 | 54.8 | 1.03 | 15.3 | 7.61 | 13.38 | 1527 | 31 | OK
Chip 10: | 123.4 | 8.05 | 55.4 | 1.13 | 13.2 | 7.22 | 13.29 | 1499 | 29 | OK
Chip 11: | 124.9 | 7.74 | 54.1 | 1.18 | 16.9 | 7.17 | 13.42 | 1551 | 32 | OK
#
#BadChannelSummary - not for the database
# at 2.00fC
# |  | lost | dodgy | dead | stuck | ineff | unbon | lo_gn | hi_gn | lo_off | hi_off | partbon | hi_nse
#Chip 0: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 1: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 2: | 2 | 2 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 2
#Chip 3: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 4: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 5: | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1
#Chip 6: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 7: | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1
#Chip 8: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 9: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 10: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 11: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Link 0: | 2 | 3 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 3
#Link 1: | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1
#Link 2: | 2 | 4 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 4
#
%DEFECTS
#name | channel | value
%NOISY | 377 | 12968.91
%NOISY | 378 | 12338.60
%UNBONDED | 379 | 582.52
%UNBONDED | 381 | 620.71
%NOISY | 681 | 1615.17
%NOISY | 917 | 1618.82
#6 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200003mb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 27/11/2001
%TEST_TIME | 21:21:25
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | NO
%PROBLEM | NO
#
%RUN | 4789
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 30.0 | deg. C
%T1 | 29.0 | deg. C
#
%TRIM_TYPE | 0
%TRIM_Q | 10.0
%N_SCANS | 16
%T_00 | 0
%T_01 | 1
%T_02 | 2
%T_03 | 3
%T_04 | 4
%T_05 | 5
%T_06 | 6
%T_07 | 7
%T_08 | 8
%T_09 | 9
%T_10 | 10
%T_11 | 11
%T_12 | 12
%T_13 | 13
%T_14 | 14
%T_15 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005mb220s30_tr0_20011127.trim
%MASK_FILE | D:\sctvar\results\0170200005mb220s30_tr0_20011127.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 0 | 87.5 | 128 | 87.0 | 1.23 | -18.4 | 1.9 | 3.88 | 0.17
Chip 1: | 0 | 87.5 | 128 | 87.0 | 1.84 | -19.8 | 2.5 | 3.55 | 0.17
Chip 2: | 0 | 87.5 | 128 | 86.9 | 1.08 | -22.3 | 2.9 | 3.03 | 0.27
Chip 3: | 0 | 87.5 | 128 | 86.9 | 1.23 | -18.6 | 1.9 | 3.82 | 0.16
Chip 4: | 0 | 87.5 | 127 | 86.9 | 1.12 | -20.5 | 2.1 | 3.55 | 0.19
Chip 5: | 0 | 87.5 | 128 | 87.1 | 1.31 | -21.1 | 2.6 | 3.39 | 0.12
Chip 6: | 0 | 87.5 | 128 | 87.0 | 1.24 | -19.9 | 2.2 | 3.58 | 0.17
Chip 7: | 0 | 87.5 | 128 | 87.0 | 1.22 | -19.9 | 2.1 | 3.55 | 0.18
Chip 8: | 0 | 87.5 | 128 | 86.8 | 1.21 | -19.9 | 2.4 | 3.48 | 0.15
Chip 9: | 0 | 87.5 | 128 | 87.0 | 1.12 | -20.2 | 2.5 | 3.37 | 0.15
Chip 10: | 0 | 87.5 | 128 | 86.8 | 1.20 | -19.2 | 2.3 | 3.47 | 0.14
Chip 11: | 0 | 87.5 | 127 | 86.9 | 1.47 | -16.9 | 1.9 | 4.13 | 0.36
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -11.50
%TR_STEP | 165 | 4.61
%TR_STEP | 257 | 3.84
%TR_OFFSET | 378 | -12.35
%TR_OFFSET | 452 | -11.54
%TR_STEP | 511 | 4.36
%TR_STEP | 639 | 4.19
%TR_OFFSET | 734 | -13.23
%TR_OFFSET | 210 | -26.78
%TR_STEP | 251 | 4.17
%TR_STEP | 255 | 4.23
%TR_OFFSET | 585 | -26.96
%TR_OFFSET | 750 | -25.06
#13 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200003mb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 27/11/2001
%TEST_TIME | 21:21:25
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 4789
%SCAN | 17
#
%HOST | ATLASPC11
%T0 | 30.0 | deg. C
%T1 | 29.0 | deg. C
#
%TRIM_TYPE | 1
%TRIM_Q | 10.0
%N_SCANS | 4
%T_00 | 3
%T_01 | 7
%T_02 | 11
%T_03 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005mb220s30_tr1_20011127.trim
%MASK_FILE | D:\sctvar\results\0170200005mb220s30_tr1_20011127.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 1 | 117.5 | 128 | 117.1 | 2.12 | -10.1 | 1.0 | 7.19 | 0.30
Chip 1: | 1 | 117.5 | 128 | 117.3 | 1.88 | -10.6 | 1.5 | 6.72 | 0.43
Chip 2: | 1 | 117.5 | 128 | 117.0 | 1.86 | -10.9 | 1.4 | 6.17 | 0.44
Chip 3: | 1 | 117.5 | 128 | 117.1 | 2.07 | -9.9 | 1.1 | 7.21 | 0.29
Chip 4: | 1 | 117.5 | 128 | 116.6 | 1.91 | -10.7 | 1.2 | 6.83 | 0.32
Chip 5: | 1 | 117.5 | 128 | 117.2 | 1.86 | -11.1 | 1.3 | 6.49 | 0.21
Chip 6: | 1 | 117.5 | 128 | 116.8 | 1.88 | -10.6 | 1.2 | 6.82 | 0.27
Chip 7: | 1 | 117.5 | 128 | 117.1 | 2.12 | -10.1 | 1.1 | 7.00 | 0.31
Chip 8: | 1 | 117.5 | 128 | 117.1 | 1.97 | -10.3 | 1.2 | 6.72 | 0.27
Chip 9: | 1 | 117.5 | 128 | 117.0 | 1.83 | -10.5 | 1.4 | 6.52 | 0.26
Chip 10: | 1 | 117.5 | 128 | 117.2 | 1.88 | -9.9 | 1.2 | 6.76 | 0.25
Chip 11: | 1 | 117.5 | 128 | 117.3 | 2.47 | -9.0 | 1.0 | 7.75 | 0.59
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -6.50
%TR_STEP | 165 | 10.64
%TR_STEP | 258 | 7.55
%TR_OFFSET | 378 | -6.07
%TR_OFFSET | 452 | -5.71
%TR_STEP | 511 | 8.19
%TR_STEP | 639 | 7.85
%TR_OFFSET | 691 | -7.08
%TR_OFFSET | 734 | -7.16
%TR_OFFSET | 171 | -14.05
%TR_OFFSET | 210 | -13.68
%TR_STEP | 251 | 8.00
%TR_STEP | 255 | 8.05
%TR_OFFSET | 585 | -13.84
%TR_STEP | 627 | 7.61
%TR_OFFSET | 750 | -13.49
%TR_STEP | 765 | 9.69
%TR_STEP | 766 | 9.60
#18 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200003mb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 27/11/2001
%TEST_TIME | 21:21:25
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 4789
%SCAN | 21
#
%HOST | ATLASPC11
%T0 | 30.0 | deg. C
%T1 | 29.0 | deg. C
#
%TRIM_TYPE | 2
%TRIM_Q | 10.0
%N_SCANS | 4
%T_00 | 3
%T_01 | 7
%T_02 | 11
%T_03 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005mb220s30_tr2_20011127.trim
%MASK_FILE | D:\sctvar\results\0170200005mb220s30_tr2_20011127.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 2 | 137.5 | 128 | 136.8 | 3.21 | -6.7 | 0.6 | 10.72 | 0.47
Chip 1: | 2 | 137.5 | 128 | 137.5 | 2.66 | -7.0 | 1.0 | 10.05 | 0.63
Chip 2: | 2 | 137.5 | 128 | 136.8 | 2.63 | -7.2 | 0.9 | 9.32 | 0.59
Chip 3: | 2 | 137.5 | 128 | 137.0 | 3.12 | -6.6 | 0.7 | 10.78 | 0.40
Chip 4: | 2 | 137.5 | 128 | 137.2 | 2.94 | -7.1 | 0.8 | 10.19 | 0.45
Chip 5: | 2 | 137.5 | 128 | 136.8 | 2.70 | -7.4 | 0.8 | 9.71 | 0.34
Chip 6: | 2 | 137.5 | 128 | 137.0 | 2.66 | -7.0 | 0.7 | 10.22 | 0.42
Chip 7: | 2 | 137.5 | 128 | 136.6 | 3.12 | -6.7 | 0.7 | 10.53 | 0.43
Chip 8: | 2 | 137.5 | 128 | 137.4 | 2.86 | -6.9 | 0.8 | 9.98 | 0.39
Chip 9: | 2 | 137.5 | 128 | 137.3 | 2.85 | -7.0 | 0.9 | 9.73 | 0.39
Chip 10: | 2 | 137.5 | 128 | 136.7 | 2.85 | -6.6 | 0.8 | 10.06 | 0.38
Chip 11: | 2 | 137.5 | 128 | 136.9 | 3.31 | -6.2 | 0.7 | 11.25 | 0.77
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -4.23
%TR_STEP | 165 | 15.74
%TR_STEP | 258 | 11.12
%TR_OFFSET | 378 | -3.81
%TR_OFFSET | 452 | -4.10
%TR_STEP | 511 | 12.03
%TR_OFFSET | 608 | -4.64
%TR_STEP | 639 | 11.70
%TR_OFFSET | 691 | -4.65
%TR_OFFSET | 734 | -4.60
%TR_OFFSET | 171 | -9.09
%TR_OFFSET | 210 | -9.13
%TR_OFFSET | 585 | -9.08
%TR_STEP | 627 | 11.30
%TR_OFFSET | 750 | -9.32
%TR_STEP | 765 | 13.73
#16 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200003mb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 27/11/2001
%TEST_TIME | 21:21:25
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 4789
%SCAN | 25
#
%HOST | ATLASPC11
%T0 | 30.0 | deg. C
%T1 | 29.0 | deg. C
#
%TRIM_TYPE | 3
%TRIM_Q | 10.0
%N_SCANS | 4
%T_00 | 3
%T_01 | 7
%T_02 | 11
%T_03 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005mb220s30_tr3_20011127.trim
%MASK_FILE | D:\sctvar\results\0170200005mb220s30_tr3_20011127.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 3 | 157.5 | 128 | 156.6 | 4.24 | -5.1 | 0.5 | 14.10 | 0.55
Chip 1: | 3 | 157.5 | 128 | 157.5 | 3.51 | -5.2 | 0.8 | 13.33 | 0.81
Chip 2: | 3 | 157.5 | 128 | 157.0 | 3.55 | -5.4 | 0.7 | 12.44 | 0.71
Chip 3: | 3 | 157.5 | 128 | 157.2 | 4.06 | -5.0 | 0.5 | 14.17 | 0.49
Chip 4: | 3 | 157.5 | 128 | 157.3 | 3.90 | -5.4 | 0.6 | 13.41 | 0.58
Chip 5: | 3 | 157.5 | 128 | 157.1 | 3.60 | -5.5 | 0.6 | 13.00 | 0.43
Chip 6: | 3 | 157.5 | 128 | 156.8 | 3.69 | -5.2 | 0.6 | 13.55 | 0.51
Chip 7: | 3 | 157.5 | 128 | 156.7 | 4.00 | -5.1 | 0.6 | 13.86 | 0.55
Chip 8: | 3 | 157.5 | 128 | 156.6 | 3.76 | -5.2 | 0.6 | 13.24 | 0.50
Chip 9: | 3 | 157.5 | 128 | 157.3 | 3.69 | -5.2 | 0.7 | 12.92 | 0.49
Chip 10: | 3 | 157.5 | 128 | 156.7 | 3.96 | -5.0 | 0.6 | 13.34 | 0.49
Chip 11: | 3 | 157.5 | 128 | 156.5 | 4.09 | -4.7 | 0.5 | 14.82 | 0.91
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -3.24
%TR_STEP | 165 | 20.79
%TR_STEP | 258 | 14.60
%TR_OFFSET | 378 | -3.04
%TR_OFFSET | 452 | -3.03
%TR_STEP | 639 | 15.19
%TR_OFFSET | 691 | -3.61
%TR_OFFSET | 734 | -3.56
%TR_OFFSET | 171 | -6.80
%TR_OFFSET | 210 | -6.98
%TR_STEP | 251 | 15.57
%TR_STEP | 351 | 14.84
%TR_OFFSET | 585 | -7.03
%TR_STEP | 627 | 14.82
%TR_OFFSET | 750 | -7.12
%TR_STEP | 765 | 17.90
#16 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200003mb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 27/11/2001
%TEST_TIME | 21:21:25
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 4789
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 30.0 | deg. C
%T1 | 29.0 | deg. C
#
%TRIM_TYPE | -1
%TRIM_Q | 10.0
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005mb220s30_tr-1_20011127.trim
%MASK_FILE | D:\sctvar\results\0170200005mb220s30_tr-1_20011127.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 0 | 100.0 | 128 | 100.0 | 1.12 | -18.4 | 1.9 | 3.88 | 0.17
Chip 1: | 0 | 100.0 | 128 | 100.1 | 1.04 | -19.8 | 2.5 | 3.55 | 0.17
Chip 2: | 1 | 100.0 | 128 | 99.9 | 1.81 | -10.9 | 1.4 | 6.17 | 0.44
Chip 3: | 1 | 100.0 | 128 | 99.7 | 2.05 | -9.9 | 1.1 | 7.21 | 0.29
Chip 4: | 0 | 100.0 | 128 | 100.1 | 1.01 | -20.5 | 2.1 | 3.55 | 0.19
Chip 5: | 0 | 100.0 | 128 | 100.1 | 1.05 | -21.1 | 2.6 | 3.39 | 0.12
Chip 6: | 0 | 100.0 | 128 | 100.0 | 1.00 | -19.9 | 2.2 | 3.58 | 0.17
Chip 7: | 0 | 100.0 | 128 | 100.0 | 1.03 | -19.9 | 2.1 | 3.55 | 0.18
Chip 8: | 0 | 100.0 | 128 | 100.0 | 0.97 | -19.9 | 2.4 | 3.48 | 0.15
Chip 9: | 1 | 100.0 | 128 | 99.7 | 1.85 | -10.5 | 1.4 | 6.52 | 0.26
Chip 10: | 0 | 100.0 | 128 | 100.0 | 1.01 | -19.2 | 2.3 | 3.47 | 0.14
Chip 11: | 1 | 100.0 | 128 | 99.6 | 2.42 | -9.0 | 1.0 | 7.75 | 0.59
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -11.50
%TR_STEP | 165 | 4.61
%TR_STEP | 258 | 7.55
%TR_OFFSET | 378 | -6.07
%TR_OFFSET | 452 | -5.71
%TR_STEP | 511 | 8.19
%TR_STEP | 639 | 4.19
%TR_OFFSET | 734 | -13.23
%TR_OFFSET | 210 | -26.78
%TR_STEP | 251 | 4.17
%TR_STEP | 255 | 4.23
%TR_OFFSET | 585 | -26.96
%TR_OFFSET | 750 | -13.49
%TR_STEP | 765 | 9.69
%TR_STEP | 766 | 9.60
#15 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30q1t087r0ini
%TEST_NAME | ResponseCurve
%TEST_DATE | 27/11/2001
%TEST_TIME | 23:17:23
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4801
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 29.0 | deg. C
%T1 | 29.0 | deg. C
#
#
%N_SCANS | 10
%Q_00 | 0.50 fC
%Q_01 | 0.75 fC
%Q_02 | 1.00 fC
%Q_03 | 1.25 fC
%Q_04 | 1.50 fC
%Q_05 | 2.00 fC
%Q_06 | 3.00 fC
%Q_07 | 4.00 fC
%Q_08 | 6.00 fC
%Q_09 | 8.00 fC
%RC_FIT | 3 | (exponential fit)
#
%ModuleResponseCurvePlot
# |  | func | p0 | p1 | p2
Chip 0: | 3 | 1717.94 | 7.44 | -826.66
Chip 1: | 3 | 1653.86 | 7.26 | -793.58
Chip 2: | 3 | 1482.04 | 6.77 | -706.17
Chip 3: | 3 | 1684.12 | 7.45 | -808.56
Chip 4: | 3 | 1683.73 | 7.31 | -809.36
Chip 5: | 3 | 1584.54 | 7.07 | -758.17
Chip 6: | 3 | 1636.59 | 7.26 | -784.61
Chip 7: | 3 | 1714.27 | 7.59 | -823.10
Chip 8: | 3 | 1604.50 | 7.18 | -767.94
Chip 9: | 3 | 1626.15 | 7.39 | -777.56
Chip 10: | 3 | 1605.67 | 7.33 | -766.56
Chip 11: | 3 | 2002.30 | 9.11 | -965.45
#
%ModuleGainOffsetNoise
# at 2.00fC
# |  | vt50 | rms | gain | rms | offset | rms | outnse | innse | rms | comment
Chip 0: | 146.2 | 1.57 | 56.6 | 1.04 | 32.5 | 1.59 | 13.28 | 1467 | 26 | OK
Chip 1: | 146.2 | 2.11 | 55.8 | 1.03 | 33.4 | 2.17 | 13.29 | 1488 | 30 | OK
Chip 2: | 144.1 | 1.71 | 53.5 | 1.16 | 34.9 | 1.61 | 12.99 | 1517 | 119 | OK
Chip 3: | 145.8 | 1.61 | 55.5 | 1.17 | 33.6 | 1.77 | 13.53 | 1525 | 29 | OK
Chip 4: | 147.2 | 1.60 | 56.5 | 1.18 | 32.7 | 1.68 | 13.78 | 1525 | 29 | OK
Chip 5: | 144.9 | 1.69 | 54.8 | 1.08 | 34.2 | 1.60 | 13.10 | 1494 | 36 | OK
Chip 6: | 145.5 | 1.57 | 55.2 | 1.11 | 33.8 | 1.69 | 13.20 | 1493 | 33 | OK
Chip 7: | 146.4 | 1.52 | 55.3 | 0.96 | 34.3 | 1.52 | 14.04 | 1587 | 30 | OK
Chip 8: | 145.0 | 1.57 | 54.7 | 1.05 | 34.4 | 1.52 | 13.59 | 1553 | 31 | OK
Chip 9: | 145.0 | 1.46 | 54.0 | 0.96 | 35.6 | 1.46 | 13.50 | 1563 | 29 | OK
Chip 10: | 145.6 | 1.51 | 53.8 | 1.08 | 36.3 | 1.50 | 13.49 | 1567 | 33 | OK
Chip 11: | 144.5 | 1.94 | 54.2 | 1.27 | 35.8 | 1.80 | 13.95 | 1607 | 39 | OK
#
#BadChannelSummary - not for the database
# at 2.00fC
# |  | lost | dodgy | dead | stuck | ineff | unbon | lo_gn | hi_gn | lo_off | hi_off | partbon | hi_nse
#Chip 0: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 1: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 2: | 4 | 0 | 2 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 3: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 4: | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 5: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 6: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 7: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 8: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 9: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 10: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 11: | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Link 0: | 5 | 0 | 3 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0
#Link 1: | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Link 2: | 6 | 0 | 4 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0
#
%DEFECTS
#name | channel | value
%DEAD | 377
%DEAD | 378
%UNBONDED | 379 | 599.17
%UNBONDED | 381 | 624.48
%DEAD | 628
%DEAD | 1518
#6 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30q1t087r0ini
%TEST_NAME | Timewalk
%TEST_DATE | 27/11/2001
%TEST_TIME | 23:59:42
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | NO
%PROBLEM | NO
#
%RUN | 4805
%SCAN | 2
#
%HOST | ATLASPC11
%T0 | 29.0 | deg. C
%T1 | 28.0 | deg. C
#
#
# run scan charge
# 4805 2 1.25
# 4805 3 1.50
# 4805 4 1.75
# 4805 5 2.00
# 4805 6 3.00
# 4805 7 4.00
# 4805 8 5.00
# 4805 9 6.00
# 4805 10 8.00
# 4805 11 10.00
#
%00_TW | 0.0
%01_TW | 0.0
%02_TW | 0.0
%03_TW | 0.0
%04_TW | 0.0
%05_TW | 0.0
%06_TW | 0.0
%07_TW | 0.0
%08_TW | 0.0
%09_TW | 0.0
%10_TW | 0.0
%11_TW | 0.0
#
%00_TCAL | 11.4
%01_TCAL | 12.3
%02_TCAL | 12.6
%03_TCAL | 11.9
%04_TCAL | 12.1
%05_TCAL | 12.2
%06_TCAL | 12.2
%07_TCAL | 11.6
%08_TCAL | 12.4
%09_TCAL | 12.6
%10_TCAL | 12.8
%11_TCAL | 13.8
#
%DEFECTS
#name | chip | timewalk
%TW_LO | 0 | 0.0
%TW_LO | 1 | 0.0
%TW_LO | 2 | 0.0
%TW_LO | 3 | 0.0
%TW_LO | 4 | 0.0
%TW_LO | 5 | 0.0
%TW_LO | 6 | 0.0
%TW_LO | 7 | 0.0
%TW_LO | 8 | 0.0
%TW_LO | 9 | 0.0
%TW_LO | 10 | 0.0
%TW_LO | 11 | 0.0
#12 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30q1t087r0ini
%TEST_NAME | NoiseOccupancy
%TEST_DATE | 28/11/2001
%TEST_TIME | 03:54:59
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4844
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 28.0 | deg. C
%T1 | 28.0 | deg. C
#
#
%IccMax: 940
%IddMax: 810
%IccTyp: 940
%IddTyp: 510
#
%NOSummary
#
# |  | Offset | MeanOcc | RMSOcc | EstENC
Chip 0: | 15.7 | 1.7e-006 | 1.7e-006 | 1669
Chip 1: | 15.9 | 2.3e-006 | 2.5e-006 | 1706
Chip 2: | -12.5 | 1.6e-005 | 1.2e-004 | 1994
Chip 3: | 12.8 | 1.1e-005 | 9.1e-006 | 1852
Chip 4: | 9.1 | 1.7e-005 | 1.3e-005 | 1898
Chip 5: | 9.7 | 3.0e-005 | 3.3e-005 | 2042
Chip 6: | 13.4 | 8.5e-006 | 8.6e-006 | 1842
Chip 7: | 12.9 | 1.7e-005 | 9.6e-006 | 1851
Chip 8: | 11.1 | 1.9e-005 | 1.2e-005 | 1932
Chip 9: | 12.2 | 2.8e-005 | 1.7e-005 | 1986
Chip 10: | 11.4 | 3.2e-005 | 1.7e-005 | 2010
Chip 11: | 4.3 | 1.1e-004 | 1.0e-004 | 2269
#
%DEFECTS
#name | chan
%NO_LO | 12 | 0.0e+000
%NO_LO | 13 | 0.0e+000
%NO_LO | 15 | 0.0e+000
%NO_LO | 21 | 0.0e+000
%NO_LO | 22 | 0.0e+000
%NO_LO | 23 | 0.0e+000
%NO_LO | 28 | 0.0e+000
%NO_LO | 34 | 0.0e+000
%NO_LO | 36 | 0.0e+000
%NO_LO | 37 | 0.0e+000
%NO_LO | 40 | 0.0e+000
%NO_LO | 42 | 0.0e+000
%NO_LO | 44 | 0.0e+000
%NO_LO | 45 | 0.0e+000
%NO_LO | 51 | 0.0e+000
%NO_LO | 55 | 0.0e+000
%NO_LO | 66 | 0.0e+000
%NO_LO | 68 | 0.0e+000
%NO_LO | 70 | 0.0e+000
%NO_LO | 74 | 0.0e+000
%NO_LO | 81 | 0.0e+000
%NO_LO | 82 | 0.0e+000
%NO_LO | 83 | 0.0e+000
%NO_LO | 88 | 0.0e+000
%NO_LO | 89 | 0.0e+000
%NO_LO | 90 | 0.0e+000
%NO_LO | 93 | 0.0e+000
%NO_LO | 97 | 0.0e+000
%NO_LO | 102 | 0.0e+000
%NO_LO | 103 | 0.0e+000
%NO_LO | 104 | 0.0e+000
%NO_LO | 105 | 0.0e+000
%NO_LO | 109 | 0.0e+000
%NO_LO | 111 | 0.0e+000
%NO_LO | 125 | 0.0e+000
%NO_LO | 140 | 0.0e+000
%NO_LO | 142 | 0.0e+000
%NO_LO | 144 | 0.0e+000
%NO_LO | 147 | 0.0e+000
%NO_LO | 152 | 0.0e+000
%NO_LO | 156 | 0.0e+000
%NO_LO | 159 | 0.0e+000
%NO_LO | 163 | 0.0e+000
%NO_LO | 165 | 0.0e+000
%NO_LO | 175 | 0.0e+000
%NO_LO | 176 | 0.0e+000
%NO_LO | 184 | 0.0e+000
%NO_LO | 186 | 0.0e+000
%NO_LO | 188 | 0.0e+000
%NO_LO | 189 | 0.0e+000
%NO_LO | 194 | 0.0e+000
%NO_LO | 201 | 0.0e+000
%NO_LO | 202 | 0.0e+000
%NO_LO | 204 | 0.0e+000
%NO_LO | 207 | 0.0e+000
%NO_LO | 212 | 0.0e+000
%NO_LO | 213 | 0.0e+000
%NO_LO | 221 | 0.0e+000
%NO_LO | 223 | 0.0e+000
%NO_LO | 224 | 0.0e+000
%NO_LO | 226 | 0.0e+000
%NO_LO | 229 | 0.0e+000
%NO_LO | 231 | 0.0e+000
%NO_LO | 249 | 0.0e+000
%NO_LO | 266 | 0.0e+000
%NO_LO | 278 | 0.0e+000
%NO_LO | 280 | 0.0e+000
%NO_LO | 282 | 0.0e+000
%NO_LO | 292 | 0.0e+000
%NO_LO | 298 | 0.0e+000
%NO_LO | 308 | 0.0e+000
%NO_LO | 310 | 0.0e+000
%NO_LO | 313 | 0.0e+000
%NO_LO | 315 | 0.0e+000
%NO_LO | 318 | 0.0e+000
%NO_LO | 350 | 0.0e+000
%NO_LO | 370 | 0.0e+000
%NO_LO | 377 | 0.0e+000
%NO_LO | 378 | 0.0e+000
%NO_LO | 379 | 0.0e+000
%NO_HI | 380 | 1.3e-003
%NO_LO | 381 | 0.0e+000
%NO_LO | 417 | 0.0e+000
%NO_LO | 424 | 0.0e+000
%NO_LO | 628 | 0.0e+000
%NO_LO | 876 | 0.0e+000
%NO_LO | 1518 | 0.0e+000
%NO_HI | 1529 | 5.0e-004
%NO_HI | 1531 | 5.2e-004
%NO_HI | 1533 | 7.8e-004
#90 defects found

## SURVEY_XY-TC
atlasPartsId | 20220170200005
eventDescription | SURVEY_XY-TC | : X=SB, INITIAL, TC, LT, ...
date [dd/mm/yyyy] | 17/7/2001
location [instituteCode(DB)] | KEK
personInitials | TKr
problem [YES/NO] | NO
pass [YES/NO] | YES
comments | any comments
temperature(xx.x) [C] | 25.6
measurementJigID | <data>
mhx [um] | 5.4205
mhy [um] | 0.94674
msx [um] | 1.6715
msy [um] | -0.22135
sepf [um] | 1.451
sepb [um] | -2.5208
midxf [um] | 0.92394
midyf [um] | -0.55795
a1 [mrad] | -0.0018193
a2 [mrad] | -0.0055839
a3 [mrad] | -0.006137
a4 [mrad] | 0.0064067
stereo [mrad] | -0.037186
hymxf [um] | 58.7146
hymyf [um] | -0.18078
hymaf [mrad] | 0.38978
hymxb [um] | 0.23568
hymyb [um] | 5.6637
hymab [mrad] | -0.71143
conp1x [um] | 85.7851
conp1y [um] | 1.2323

## SURVEY_Z-TC
atlasPartsId | 20220170200005
eventDescription | SURVEY_Z-TC | : X=INITIAL, TC, LT, ...
date [dd/mm/yyy] | 6/12/2001
location [instituteCode(DB)] | KEK
personInitial | TKr
problem [YES/NO] | NO
pass [YES/NO] | YES
comments | V2.2.0
temperature [C] | 26.5
measurementJigID | <data>
maxZlower [mm] | -0.034636
maxZupper [mm] | 0.0061953
midplaneEq | z=ax+by+c
Left a | 0.0004049
b | 0.00041618
c | 0.47446
Right a | -0.00078137
b | 0.00015775
c | 0.46981
midplaneHeight [mm] | 0.47214
moduleThickness [mm] | 1.1811
optimalMaxZerrorLower [mm] | 0.033573
optimalMaxZerrorUpper [mm] | 0.037488
optimalRmsZerrorLower [mm] | 0.010133
optimalRmsZerrorUpper [mm] | 0.0091692
moduleConcavity x [mm] | 0.037602
y | -0.0030906
sensorSkew x [mm] | -0.025447
y | 0.018714
coolingTabThickness [mm] | 0.93725
farTabThickness [mm] | 0.96164
halfTabThickness [mm] | 0.47472
tabSkew y [mm] | -0.012192
adhesiveThicknessTotal [mm] | 0.16168
adhesiveAsymmetry [mm] | 0.0025836
loCoolingFacing a [mrad] | -0.0081213
b [mrad] | -0.14643
loCoolingFacingConcavity [mm] | 0
optimalFrameRawDATA
leftLower | x
1 | -63.1985
2 | -47.3269
3 | -31.4554
4 | -15.5839
5 | 0.28762
6 | -63.5159
7 | -47.6443
8 | -31.7728
9 | -15.9013
10 | -0.02972
11 | -63.8333
12 | -47.9617
13 | -32.0902
14 | -16.2186
15 | -0.34718
16 | -64.1507
17 | -48.2791
18 | -32.4075
19 | -16.5361
20 | -0.66452
21 | -64.4682
22 | -48.5965
23 | -32.7249
24 | -16.8535
25 | -0.98201
rightLower | x
1 | 0.98153
2 | 16.8547
3 | 32.7247
4 | 48.5961
5 | 64.4677
6 | 0.6641
7 | 16.5383
8 | 32.4072
9 | 48.2787
10 | 64.1502
11 | 0.34674
12 | 16.2199
13 | 32.0898
14 | 47.9613
15 | 63.8328
16 | 0.029387
17 | 15.9035
18 | 31.7725
19 | 47.644
20 | 63.5154
21 | -0.2881
22 | 15.5851
23 | 31.455
24 | 47.3265
25 | 63.1979
leftUpper | x
1 | -63.1512
2 | -47.2796
3 | -31.408
4 | -15.5364
5 | 0.33528
6 | -63.4669
7 | -47.5952
8 | -31.7236
9 | -15.8521
10 | 0.019666
11 | -63.7825
12 | -47.911
13 | -32.0392
14 | -16.1677
15 | -0.29603
16 | -64.0981
17 | -48.2265
18 | -32.3549
19 | -16.4832
20 | -0.61153
21 | -64.4137
22 | -48.5422
23 | -32.6705
24 | -16.7989
25 | -0.92722
rightUpper | x
1 | 0.92752
2 | 16.7979
3 | 32.6708
4 | 48.5425
5 | 64.4143
6 | 0.61182
7 | 16.4816
8 | 32.3552
9 | 48.2268
10 | 64.0986
11 | 0.29613
12 | 16.1665
13 | 32.0395
14 | 47.9113
15 | 63.7829
16 | -0.019378
17 | 15.8503
18 | 31.7239
19 | 47.5957
20 | 63.4673
21 | -0.33517
22 | 15.5354
23 | 31.4082
24 | 47.28
25 | 63.1516
leftLower | y
1 | -32.8277
2 | -32.5081
3 | -32.1883
4 | -31.8684
5 | -31.5488
6 | -17.0563
7 | -16.7366
8 | -16.4169
9 | -16.097
10 | -15.7773
11 | -1.2848
12 | -0.96509
13 | -0.64546
14 | -0.32561
15 | -0.0057501
16 | 14.4867
17 | 14.8064
18 | 15.126
19 | 15.4458
20 | 15.7655
21 | 30.2582
22 | 30.5778
23 | 30.8976
24 | 31.2173
25 | 31.5371
rightLower | y
1 | -31.5364
2 | -31.2189
3 | -30.897
4 | -30.5773
5 | -30.2574
6 | -15.7651
7 | -15.4487
8 | -15.1257
9 | -14.8058
10 | -14.486
11 | 0.0062769
12 | 0.32399
13 | 0.64578
14 | 0.96552
15 | 1.2854
16 | 15.7778
17 | 16.0943
18 | 16.4172
19 | 16.7371
20 | 17.0568
21 | 31.5492
22 | 31.8668
23 | 32.1887
24 | 32.5085
25 | 32.8284
leftUpper | y
1 | -32.8216
2 | -32.5037
3 | -32.186
4 | -31.8683
5 | -31.5506
6 | -17.0506
7 | -16.7328
8 | -16.4151
9 | -16.0974
10 | -15.7798
11 | -1.2797
12 | -0.96196
13 | -0.64426
14 | -0.3266
15 | -0.0089385
16 | 14.4911
17 | 14.8088
18 | 15.1266
19 | 15.4442
20 | 15.7619
21 | 30.262
22 | 30.5796
23 | 30.8974
24 | 31.2151
25 | 31.5327
rightUpper | y
1 | -31.533
2 | -31.2132
3 | -30.8978
4 | -30.5803
5 | -30.2626
6 | -15.7623
7 | -15.4414
8 | -15.127
9 | -14.8094
10 | -14.4919
11 | 0.0085441
12 | 0.32836
13 | 0.64382
14 | 0.96133
15 | 1.279
16 | 15.7794
17 | 16.1003
18 | 16.4146
19 | 16.7322
20 | 17.0498
21 | 31.5502
22 | 31.87
23 | 32.1855
24 | 32.5031
25 | 32.8206
leftLower | z
1 | 0.087416
2 | 0.040582
3 | 0.022029
4 | 0.016994
5 | 0.014371
6 | 0.060259
7 | 0.016758
8 | 0.0047812
9 | 0.0022319
10 | 0.0015126
11 | 0.053543
12 | 0.012463
13 | 0.0020944
14 | -0.0017679
15 | 0.0028268
16 | 0.062503
17 | 0.016463
18 | 0.00024958
19 | -0.0059177
20 | -0.0085421
21 | 0.08935
22 | 0.038755
23 | 0.014484
24 | 0.0056515
25 | 0.0055542
rightLower | z
1 | 0.0098044
2 | 0.012781
3 | 0.015758
4 | 0.040714
5 | 0.089714
6 | -0.00072726
7 | -0.0016964
8 | -0.0026669
9 | 0.014601
10 | 0.058691
11 | 0.0024592
12 | -0.0021087
13 | -0.0066775
14 | 0.0050127
15 | 0.046617
16 | 8.6677e-05
17 | -0.0023211
18 | -0.0047299
19 | 0.0093873
20 | 0.051509
21 | 0.017303
22 | 0.014068
23 | 0.010832
24 | 0.031945
25 | 0.078338
leftUpper | z
1 | -0.097287
2 | -0.046485
3 | -0.018567
4 | -0.0079742
5 | -0.009118
6 | -0.067228
7 | -0.018699
8 | -0.0017871
9 | 0.0028682
10 | -0.00087064
11 | -0.055273
12 | -0.010429
13 | 0.0016823
14 | 0.00084955
15 | -0.0056842
16 | -0.059976
17 | -0.013359
18 | 0.0013007
19 | 0.0017949
20 | -0.0024419
21 | -0.083258
22 | -0.0348
23 | -0.011856
24 | -0.0074877
25 | -0.010556
rightUpper | z
1 | -0.013039
2 | -0.013975
3 | -0.014911
4 | -0.030139
5 | -0.070242
6 | -0.0012599
7 | -0.0032531
8 | -0.0052462
9 | -0.015214
10 | -0.05292
11 | -0.0065013
12 | -0.004426
13 | -0.0023506
14 | -0.011969
15 | -0.050022
16 | 0.0042624
17 | 0.0012876
18 | -0.0016871
19 | -0.01677
20 | -0.057863
21 | 0.0011818
22 | -0.0034033
23 | -0.0079885
24 | -0.033357
25 | -0.078881

## ELECTRICAL-TC
atlasPartsId | 20220170200005
eventDescription | ELECTRICAL-TC |  | : X=INITIAL, TC, LT, ANNEAL, ...
date [dd/mm/yyyy] | 30/11/2001
location [instituteCode(DB)] | KEK
personInitial | ta
problem [YES/NO] | no
pass [YES/NO] | yes
comments
#
# | http://hepwww.rl.ac.uk/atlas-sct/documents/dev/Electrical_Tests_301.doc
# | N.B. Some revisions have been incorporated here, but not yet in the reference,
# | namely that CONFIGURATION will be stored as an ASCII file and
# | and that the RUN field will be used to store both RUN and SCAN number.
#
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30
%TEST_NAME | FullBypassTest
%TEST_DATE | 30/11/2001
%TEST_TIME | 15:59:31
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4971
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 27.0 | deg. C
%T1 | 26.0 | deg. C
#
#NConfigs 37
#Vdd = 4.2 - 3.3
#
%M0->S1: | 3.3 V | (Minimal tested)
%M0->S2: | 3.3 V | (Minimal tested)
%S1->S2: | 3.3 V | (Minimal tested)
%S1->S3: | 3.3 V | (Minimal tested)
%S2->S3: | 3.3 V | (Minimal tested)
%S2->S4: | 3.3 V | (Minimal tested)
%S3->S4: | 3.3 V | (Minimal tested)
%S3->E5: | 3.3 V | (Minimal tested)
%S4->E5: | 3.3 V | (Minimal tested)
%E5->S9: | 3.3 V | (Minimal tested)
%M8->S9: | 3.3 V | (Minimal tested)
%M8->S10: | 3.3 V | (Minimal tested)
%S9->S10: | 3.3 V | (Minimal tested)
%S9->S11: | 3.3 V | (Minimal tested)
%S10->S11: | 3.3 V | (Minimal tested)
%S10->S12: | 3.3 V | (Minimal tested)
%S11->S12: | 3.3 V | (Minimal tested)
%S11->E13: | 3.3 V | (Minimal tested)
%S12->E13: | 3.3 V | (Minimal tested)
%E13->S1: | 3.3 V | (Minimal tested)
#
%DEFECTS
#name | bypass_link | Vdd_min
#No defects found!
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30
%TEST_NAME | PipelineTest
%TEST_DATE | 30/11/2001
%TEST_TIME | 16:00:01
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4971
%SCAN | 11
#
%HOST | ATLASPC11
%T0 | 27.0 | deg. C
%T1 | 27.0 | deg. C
#
#
%00_GOOD | 128
%01_GOOD | 128
%02_GOOD | 128
%03_GOOD | 128
%04_GOOD | 128
%05_GOOD | 128
%06_GOOD | 128
%07_GOOD | 128
%08_GOOD | 128
%09_GOOD | 128
%10_GOOD | 128
%11_GOOD | 128
#
%DEFECTS
#name | chan | cell
#No defects found!
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30
%TEST_NAME | StrobeDelay
%TEST_DATE | 30/11/2001
%TEST_TIME | 16:04:44
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4972
%SCAN | 2
#
%HOST | ATLASPC11
%T0 | 27.0 | deg. C
%T1 | 27.0 | deg. C
#
#
#parameter | value
%00_DELAY | 11
%01_DELAY | 12
%02_DELAY | 12
%03_DELAY | 12
%04_DELAY | 12
%05_DELAY | 12
%06_DELAY | 12
%07_DELAY | 11
%08_DELAY | 12
%09_DELAY | 12
%10_DELAY | 12
%11_DELAY | 14
#
%DEFECTS
#name | chip
#No defects found!
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30
%TEST_NAME | ResponseCurve
%TEST_DATE | 30/11/2001
%TEST_TIME | 16:16:34
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4973
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 26.0 | deg. C
%T1 | 26.0 | deg. C
#
#
%N_SCANS | 3
%Q_00 | 1.50 fC
%Q_01 | 2.00 fC
%Q_02 | 2.50 fC
%RC_FIT | 4 | (linear fit)
#
%ModuleResponseCurvePlot
# |  | func | p0 | p1 | p2
Chip 0: | 4 | 21.55 | 55.57 | 0.00
Chip 1: | 4 | 19.60 | 55.55 | 0.00
Chip 2: | 4 | 15.68 | 54.25 | 0.00
Chip 3: | 4 | 20.12 | 55.90 | 0.00
Chip 4: | 4 | 19.80 | 57.21 | 0.00
Chip 5: | 4 | 23.39 | 54.13 | 0.00
Chip 6: | 4 | 21.04 | 55.08 | 0.00
Chip 7: | 4 | 17.76 | 56.51 | 0.00
Chip 8: | 4 | 18.49 | 54.98 | 0.00
Chip 9: | 4 | 17.71 | 54.52 | 0.00
Chip 10: | 4 | 15.65 | 55.31 | 0.00
Chip 11: | 4 | 20.79 | 53.89 | 0.00
#
%ModuleGainOffsetNoise
# at 2.00fC
# |  | vt50 | rms | gain | rms | offset | rms | outnse | innse | rms | comment
Chip 0: | 132.3 | 7.63 | 55.6 | 1.01 | 21.9 | 7.26 | 13.11 | 1475 | 32 | OK
Chip 1: | 130.6 | 7.61 | 55.5 | 1.04 | 19.6 | 7.84 | 13.10 | 1474 | 28 | OK
Chip 2: | 125.6 | 7.01 | 54.2 | 3.70 | 15.8 | 6.45 | 13.40 | 1460 | 110 | OK
Chip 3: | 132.0 | 7.60 | 55.9 | 1.12 | 20.3 | 7.04 | 13.52 | 1512 | 27 | OK
Chip 4: | 134.4 | 8.06 | 57.2 | 1.16 | 20.0 | 7.15 | 13.48 | 1473 | 31 | OK
Chip 5: | 131.5 | 8.44 | 54.1 | 0.99 | 23.8 | 7.79 | 13.00 | 1501 | 37 | OK
Chip 6: | 131.2 | 7.62 | 55.1 | 1.03 | 21.0 | 7.53 | 12.91 | 1465 | 30 | OK
Chip 7: | 130.9 | 6.55 | 56.5 | 1.09 | 17.8 | 6.68 | 13.58 | 1502 | 36 | OK
Chip 8: | 128.4 | 7.88 | 55.0 | 1.00 | 18.5 | 7.44 | 13.21 | 1502 | 29 | OK
Chip 9: | 127.0 | 8.42 | 54.5 | 0.95 | 18.2 | 7.84 | 13.23 | 1517 | 31 | OK
Chip 10: | 126.2 | 8.11 | 55.3 | 1.10 | 15.9 | 7.59 | 13.12 | 1483 | 27 | OK
Chip 11: | 128.3 | 7.83 | 53.9 | 1.12 | 20.8 | 7.43 | 13.45 | 1560 | 37 | OK
#
#BadChannelSummary - not for the database
# at 2.00fC
# |  | lost | dodgy | dead | stuck | ineff | unbon | lo_gn | hi_gn | lo_off | hi_off | partbon | hi_nse
#Chip 0: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 1: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 2: | 2 | 2 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 2
#Chip 3: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 4: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 5: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 6: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 7: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 8: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 9: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 10: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 11: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Link 0: | 2 | 2 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 2
#Link 1: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Link 2: | 2 | 2 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 2
#
%DEFECTS
#name | channel | value
%NOISY | 377 | 13143.16
%NOISY | 378 | 12257.70
%UNBONDED | 379 | 601.93
%UNBONDED | 381 | 645.47
#4 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200003mb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 30/11/2001
%TEST_TIME | 17:00:58
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | NO
%PROBLEM | NO
#
%RUN | 4974
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 29.0 | deg. C
%T1 | 28.0 | deg. C
#
%TRIM_TYPE | 0
%TRIM_Q | 10.0
%N_SCANS | 16
%T_00 | 0
%T_01 | 1
%T_02 | 2
%T_03 | 3
%T_04 | 4
%T_05 | 5
%T_06 | 6
%T_07 | 7
%T_08 | 8
%T_09 | 9
%T_10 | 10
%T_11 | 11
%T_12 | 12
%T_13 | 13
%T_14 | 14
%T_15 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005mb220s30_tr0_20011130.trim
%MASK_FILE | D:\sctvar\results\0170200005mb220s30_tr0_20011130.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 0 | 95.0 | 128 | 95.2 | 1.34 | -19.4 | 2.0 | 3.86 | 0.18
Chip 1: | 0 | 95.0 | 128 | 95.1 | 1.31 | -20.6 | 2.6 | 3.55 | 0.17
Chip 2: | 0 | 95.0 | 127 | 95.0 | 1.09 | -23.0 | 3.1 | 3.04 | 0.26
Chip 3: | 0 | 95.0 | 128 | 94.8 | 1.16 | -19.5 | 2.0 | 3.81 | 0.16
Chip 4: | 0 | 95.0 | 127 | 94.9 | 1.27 | -21.4 | 2.3 | 3.54 | 0.19
Chip 5: | 0 | 95.0 | 128 | 95.1 | 1.22 | -22.1 | 2.7 | 3.38 | 0.12
Chip 6: | 0 | 95.0 | 128 | 95.1 | 1.26 | -20.7 | 2.3 | 3.57 | 0.16
Chip 7: | 0 | 95.0 | 128 | 94.7 | 1.17 | -20.7 | 2.2 | 3.55 | 0.18
Chip 8: | 0 | 95.0 | 128 | 95.1 | 1.05 | -20.6 | 2.5 | 3.48 | 0.15
Chip 9: | 0 | 95.0 | 128 | 95.1 | 1.22 | -21.0 | 2.7 | 3.37 | 0.15
Chip 10: | 0 | 95.0 | 128 | 95.2 | 1.18 | -20.0 | 2.4 | 3.47 | 0.14
Chip 11: | 0 | 95.0 | 127 | 95.2 | 1.43 | -17.6 | 1.9 | 4.12 | 0.36
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -12.00
%TR_STEP | 127 | 4.42
%TR_STEP | 165 | 4.56
%TR_OFFSET | 378 | -12.00
%TR_OFFSET | 452 | -12.01
%TR_STEP | 511 | 4.42
%TR_STEP | 639 | 4.15
%TR_OFFSET | 691 | -13.75
%TR_OFFSET | 160 | -27.42
%TR_OFFSET | 171 | -27.72
%TR_OFFSET | 210 | -27.34
%TR_STEP | 251 | 4.15
%TR_STEP | 255 | 4.16
%TR_OFFSET | 585 | -28.18
%TR_OFFSET | 750 | -25.02
%TR_STEP | 765 | 5.27
#16 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200003mb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 30/11/2001
%TEST_TIME | 17:00:58
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 4974
%SCAN | 17
#
%HOST | ATLASPC11
%T0 | 29.0 | deg. C
%T1 | 28.0 | deg. C
#
%TRIM_TYPE | 1
%TRIM_Q | 10.0
%N_SCANS | 4
%T_00 | 3
%T_01 | 7
%T_02 | 11
%T_03 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005mb220s30_tr1_20011130.trim
%MASK_FILE | D:\sctvar\results\0170200005mb220s30_tr1_20011130.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 1 | 120.0 | 128 | 120.2 | 2.12 | -10.5 | 1.0 | 7.18 | 0.30
Chip 1: | 1 | 120.0 | 128 | 120.1 | 1.89 | -10.9 | 1.5 | 6.73 | 0.41
Chip 2: | 1 | 120.0 | 128 | 120.0 | 1.90 | -11.2 | 1.4 | 6.19 | 0.43
Chip 3: | 1 | 120.0 | 128 | 120.0 | 2.06 | -10.3 | 1.1 | 7.21 | 0.27
Chip 4: | 1 | 120.0 | 128 | 119.7 | 1.87 | -11.1 | 1.3 | 6.81 | 0.30
Chip 5: | 1 | 120.0 | 128 | 120.0 | 1.86 | -11.7 | 1.4 | 6.46 | 0.21
Chip 6: | 1 | 120.0 | 128 | 119.6 | 1.86 | -10.9 | 1.2 | 6.83 | 0.27
Chip 7: | 1 | 120.0 | 128 | 120.2 | 2.01 | -10.5 | 1.2 | 6.99 | 0.31
Chip 8: | 1 | 120.0 | 128 | 120.0 | 1.96 | -10.7 | 1.3 | 6.71 | 0.27
Chip 9: | 1 | 120.0 | 128 | 120.1 | 1.87 | -10.9 | 1.4 | 6.51 | 0.25
Chip 10: | 1 | 120.0 | 128 | 120.2 | 1.82 | -10.2 | 1.3 | 6.77 | 0.26
Chip 11: | 1 | 120.0 | 128 | 120.2 | 2.48 | -9.5 | 1.0 | 7.72 | 0.59
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -6.39
%TR_STEP | 127 | 8.11
%TR_STEP | 165 | 10.53
%TR_OFFSET | 377 | -6.84
%TR_OFFSET | 378 | -5.97
%TR_OFFSET | 452 | -6.19
%TR_STEP | 511 | 8.11
%TR_STEP | 639 | 7.85
%TR_OFFSET | 691 | -7.13
%TR_OFFSET | 734 | -7.17
%TR_OFFSET | 171 | -14.10
%TR_OFFSET | 210 | -14.76
%TR_STEP | 251 | 7.94
%TR_STEP | 351 | 7.64
%TR_OFFSET | 585 | -14.63
%TR_OFFSET | 750 | -13.62
%TR_STEP | 765 | 9.61
#17 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200003mb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 30/11/2001
%TEST_TIME | 17:00:58
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 4974
%SCAN | 21
#
%HOST | ATLASPC11
%T0 | 29.0 | deg. C
%T1 | 28.0 | deg. C
#
%TRIM_TYPE | 2
%TRIM_Q | 10.0
%N_SCANS | 4
%T_00 | 3
%T_01 | 7
%T_02 | 11
%T_03 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005mb220s30_tr2_20011130.trim
%MASK_FILE | D:\sctvar\results\0170200005mb220s30_tr2_20011130.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 2 | 140.0 | 128 | 139.6 | 3.29 | -7.0 | 0.6 | 10.70 | 0.47
Chip 1: | 2 | 140.0 | 128 | 140.5 | 2.66 | -7.3 | 1.0 | 10.03 | 0.62
Chip 2: | 2 | 140.0 | 128 | 139.7 | 2.73 | -7.4 | 0.9 | 9.32 | 0.59
Chip 3: | 2 | 140.0 | 128 | 139.9 | 3.25 | -6.9 | 0.7 | 10.76 | 0.40
Chip 4: | 2 | 140.0 | 128 | 140.3 | 3.01 | -7.5 | 0.8 | 10.17 | 0.44
Chip 5: | 2 | 140.0 | 128 | 139.9 | 2.75 | -7.8 | 0.8 | 9.69 | 0.34
Chip 6: | 2 | 140.0 | 128 | 139.7 | 2.74 | -7.3 | 0.7 | 10.20 | 0.41
Chip 7: | 2 | 140.0 | 128 | 140.0 | 3.14 | -6.9 | 0.8 | 10.51 | 0.43
Chip 8: | 2 | 140.0 | 128 | 140.1 | 2.89 | -7.2 | 0.8 | 9.98 | 0.39
Chip 9: | 2 | 140.0 | 128 | 140.2 | 2.85 | -7.3 | 0.9 | 9.73 | 0.37
Chip 10: | 2 | 140.0 | 128 | 139.7 | 2.86 | -6.9 | 0.8 | 10.03 | 0.38
Chip 11: | 2 | 140.0 | 128 | 139.7 | 3.34 | -6.5 | 0.7 | 11.22 | 0.77
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -4.53
%TR_STEP | 165 | 15.68
%TR_OFFSET | 378 | -4.27
%TR_OFFSET | 452 | -4.38
%TR_STEP | 511 | 12.02
%TR_STEP | 639 | 11.49
%TR_OFFSET | 691 | -4.92
%TR_OFFSET | 734 | -5.12
%TR_OFFSET | 171 | -9.23
%TR_OFFSET | 210 | -9.46
%TR_STEP | 351 | 11.20
%TR_OFFSET | 585 | -9.67
%TR_OFFSET | 750 | -9.73
%TR_STEP | 765 | 13.82
#14 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200003mb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 30/11/2001
%TEST_TIME | 17:00:58
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 4974
%SCAN | 25
#
%HOST | ATLASPC11
%T0 | 29.0 | deg. C
%T1 | 28.0 | deg. C
#
%TRIM_TYPE | 3
%TRIM_Q | 10.0
%N_SCANS | 4
%T_00 | 3
%T_01 | 7
%T_02 | 11
%T_03 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005mb220s30_tr3_20011130.trim
%MASK_FILE | D:\sctvar\results\0170200005mb220s30_tr3_20011130.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 3 | 160.0 | 128 | 159.7 | 4.22 | -5.2 | 0.5 | 14.13 | 0.55
Chip 1: | 3 | 160.0 | 128 | 160.3 | 3.60 | -5.4 | 0.8 | 13.35 | 0.80
Chip 2: | 3 | 160.0 | 128 | 160.1 | 3.63 | -5.6 | 0.7 | 12.42 | 0.71
Chip 3: | 3 | 160.0 | 128 | 160.2 | 4.11 | -5.2 | 0.5 | 14.18 | 0.50
Chip 4: | 3 | 160.0 | 128 | 160.5 | 3.86 | -5.7 | 0.6 | 13.40 | 0.58
Chip 5: | 3 | 160.0 | 128 | 159.9 | 3.66 | -5.7 | 0.6 | 12.99 | 0.44
Chip 6: | 3 | 160.0 | 128 | 160.1 | 3.73 | -5.4 | 0.5 | 13.56 | 0.52
Chip 7: | 3 | 160.0 | 128 | 159.9 | 4.08 | -5.3 | 0.6 | 13.86 | 0.55
Chip 8: | 3 | 160.0 | 128 | 160.0 | 3.74 | -5.4 | 0.6 | 13.24 | 0.51
Chip 9: | 3 | 160.0 | 128 | 160.3 | 3.66 | -5.4 | 0.6 | 12.93 | 0.49
Chip 10: | 3 | 160.0 | 128 | 159.9 | 4.02 | -5.2 | 0.6 | 13.32 | 0.50
Chip 11: | 3 | 160.0 | 128 | 159.5 | 4.14 | -4.9 | 0.5 | 14.82 | 0.93
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -3.40
%TR_STEP | 165 | 20.75
%TR_OFFSET | 377 | -3.44
%TR_OFFSET | 378 | -2.94
%TR_OFFSET | 452 | -3.28
%TR_STEP | 639 | 15.14
%TR_OFFSET | 734 | -3.71
%TR_OFFSET | 171 | -7.07
%TR_OFFSET | 210 | -7.09
%TR_STEP | 255 | 15.52
%TR_OFFSET | 585 | -7.19
%TR_STEP | 623 | 14.83
%TR_OFFSET | 750 | -7.21
%TR_STEP | 765 | 17.85
#14 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200003mb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 30/11/2001
%TEST_TIME | 17:00:58
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 4974
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 29.0 | deg. C
%T1 | 28.0 | deg. C
#
%TRIM_TYPE | -1
%TRIM_Q | 10.0
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005mb220s30_tr-1_20011130.trim
%MASK_FILE | D:\sctvar\results\0170200005mb220s30_tr-1_20011130.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 0 | 100.0 | 128 | 100.0 | 1.09 | -19.4 | 2.0 | 3.86 | 0.18
Chip 1: | 0 | 100.0 | 128 | 100.0 | 1.01 | -20.6 | 2.6 | 3.55 | 0.17
Chip 2: | 1 | 100.0 | 128 | 99.9 | 1.73 | -11.2 | 1.4 | 6.19 | 0.43
Chip 3: | 0 | 100.0 | 128 | 99.9 | 1.16 | -19.5 | 2.0 | 3.81 | 0.16
Chip 4: | 0 | 100.0 | 128 | 100.0 | 1.01 | -21.4 | 2.3 | 3.54 | 0.19
Chip 5: | 0 | 100.0 | 128 | 100.2 | 0.99 | -22.1 | 2.7 | 3.38 | 0.12
Chip 6: | 0 | 100.0 | 128 | 99.9 | 1.01 | -20.7 | 2.3 | 3.57 | 0.16
Chip 7: | 0 | 100.0 | 128 | 100.1 | 1.05 | -20.7 | 2.2 | 3.55 | 0.18
Chip 8: | 0 | 100.0 | 128 | 100.0 | 1.01 | -20.6 | 2.5 | 3.48 | 0.15
Chip 9: | 0 | 100.0 | 128 | 100.0 | 0.95 | -21.0 | 2.7 | 3.37 | 0.15
Chip 10: | 0 | 100.0 | 128 | 100.0 | 1.02 | -20.0 | 2.4 | 3.47 | 0.14
Chip 11: | 1 | 100.0 | 128 | 100.4 | 2.33 | -9.5 | 1.0 | 7.72 | 0.59
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -12.00
%TR_STEP | 127 | 4.42
%TR_STEP | 165 | 4.56
%TR_OFFSET | 377 | -6.84
%TR_OFFSET | 378 | -5.97
%TR_OFFSET | 452 | -12.01
%TR_STEP | 511 | 4.42
%TR_STEP | 639 | 4.15
%TR_OFFSET | 691 | -13.75
%TR_OFFSET | 160 | -27.42
%TR_OFFSET | 171 | -27.72
%TR_OFFSET | 210 | -27.34
%TR_STEP | 251 | 4.15
%TR_STEP | 255 | 4.16
%TR_OFFSET | 585 | -28.18
%TR_OFFSET | 750 | -13.62
%TR_STEP | 765 | 9.61
#17 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30q1t095r0TC
%TEST_NAME | ResponseCurve
%TEST_DATE | 30/11/2001
%TEST_TIME | 17:56:58
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4976
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 28.0 | deg. C
%T1 | 28.0 | deg. C
#
#
%N_SCANS | 10
%Q_00 | 0.50 fC
%Q_01 | 0.75 fC
%Q_02 | 1.00 fC
%Q_03 | 1.25 fC
%Q_04 | 1.50 fC
%Q_05 | 2.00 fC
%Q_06 | 3.00 fC
%Q_07 | 4.00 fC
%Q_08 | 6.00 fC
%Q_09 | 8.00 fC
%RC_FIT | 3 | (exponential fit)
#
%ModuleResponseCurvePlot
# |  | func | p0 | p1 | p2
Chip 0: | 3 | 1646.78 | 7.10 | -786.14
Chip 1: | 3 | 1598.73 | 7.02 | -760.83
Chip 2: | 3 | 1414.83 | 6.42 | -667.21
Chip 3: | 3 | 1589.73 | 6.93 | -756.92
Chip 4: | 3 | 1575.87 | 6.78 | -750.37
Chip 5: | 3 | 1509.22 | 6.69 | -715.56
Chip 6: | 3 | 1554.76 | 6.85 | -738.54
Chip 7: | 3 | 1620.45 | 7.12 | -771.59
Chip 8: | 3 | 1513.26 | 6.72 | -717.36
Chip 9: | 3 | 1526.02 | 6.88 | -722.91
Chip 10: | 3 | 1486.24 | 6.72 | -702.37
Chip 11: | 3 | 1773.92 | 7.96 | -847.33
#
%ModuleGainOffsetNoise
# at 2.00fC
# |  | vt50 | rms | gain | rms | offset | rms | outnse | innse | rms | comment
Chip 0: | 151.6 | 1.66 | 56.7 | 1.06 | 37.4 | 1.66 | 12.71 | 1401 | 28 | OK
Chip 1: | 151.3 | 1.42 | 55.8 | 1.05 | 38.5 | 1.78 | 12.60 | 1412 | 28 | OK
Chip 2: | 150.2 | 1.46 | 53.6 | 1.11 | 40.5 | 1.59 | 12.43 | 1449 | 110 | OK
Chip 3: | 151.7 | 1.56 | 56.1 | 1.20 | 38.0 | 1.96 | 13.06 | 1455 | 30 | OK
Chip 4: | 153.3 | 1.66 | 56.8 | 1.18 | 37.8 | 1.83 | 13.11 | 1444 | 29 | OK
Chip 5: | 150.5 | 1.45 | 54.9 | 1.09 | 39.4 | 1.66 | 12.46 | 1417 | 33 | OK
Chip 6: | 151.3 | 1.53 | 55.4 | 1.07 | 39.1 | 1.57 | 12.39 | 1398 | 28 | OK
Chip 7: | 152.0 | 1.48 | 55.7 | 0.96 | 38.9 | 1.49 | 13.24 | 1487 | 31 | OK
Chip 8: | 150.8 | 1.41 | 54.9 | 1.03 | 39.5 | 1.57 | 12.84 | 1462 | 31 | OK
Chip 9: | 150.3 | 1.49 | 54.2 | 0.94 | 40.2 | 1.42 | 12.68 | 1463 | 28 | OK
Chip 10: | 151.0 | 1.36 | 54.1 | 1.11 | 40.9 | 1.76 | 12.73 | 1470 | 30 | OK
Chip 11: | 149.9 | 1.93 | 54.7 | 1.28 | 39.8 | 1.86 | 12.95 | 1479 | 37 | OK
#
#BadChannelSummary - not for the database
# at 2.00fC
# |  | lost | dodgy | dead | stuck | ineff | unbon | lo_gn | hi_gn | lo_off | hi_off | partbon | hi_nse
#Chip 0: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 1: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 2: | 4 | 0 | 2 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 3: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 4: | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 5: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 6: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 7: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 8: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 9: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 10: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 11: | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Link 0: | 5 | 0 | 3 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0
#Link 1: | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Link 2: | 6 | 0 | 4 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0
#
%DEFECTS
#name | channel | value
%DEAD | 377
%DEAD | 378
%UNBONDED | 379 | 611.84
%UNBONDED | 381 | 623.88
%DEAD | 628
%DEAD | 1518
#6 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30q1t095r0TC
%TEST_NAME | NoiseOccupancy
%TEST_DATE | 30/11/2001
%TEST_TIME | 18:34:18
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4977
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 28.0 | deg. C
%T1 | 27.0 | deg. C
#
#
%IccMax: 940
%IddMax: 760
%IccTyp: 940
%IddTyp: 490
#
%NOSummary
#
# |  | Offset | MeanOcc | RMSOcc | EstENC
Chip 0: | 18.7 | 6.8e-006 | 4.6e-006 | 1652
Chip 1: | 18.3 | 1.0e-005 | 7.4e-006 | 1682
Chip 2: | 0.4 | 3.8e-005 | 1.9e-004 | 1931
Chip 3: | 14.9 | 3.3e-005 | 2.5e-005 | 1831
Chip 4: | 9.2 | 5.5e-005 | 3.2e-005 | 1891
Chip 5: | 11.0 | 7.9e-005 | 7.1e-005 | 2016
Chip 6: | 17.9 | 1.3e-005 | 1.5e-005 | 1709
Chip 7: | 16.0 | 2.2e-005 | 1.2e-005 | 1725
Chip 8: | 14.1 | 2.5e-005 | 1.6e-005 | 1802
Chip 9: | 15.3 | 3.6e-005 | 1.9e-005 | 1835
Chip 10: | 15.4 | 4.1e-005 | 1.9e-005 | 1850
Chip 11: | 5.3 | 1.2e-004 | 8.3e-005 | 2119
#
%DEFECTS
#name | chan
%NO_LO | 14 | 0.0e+000
%NO_LO | 42 | 0.0e+000
%NO_LO | 51 | 0.0e+000
%NO_LO | 55 | 0.0e+000
%NO_LO | 70 | 0.0e+000
%NO_LO | 74 | 0.0e+000
%NO_LO | 75 | 0.0e+000
%NO_LO | 92 | 0.0e+000
%NO_LO | 107 | 0.0e+000
%NO_LO | 108 | 0.0e+000
%NO_LO | 115 | 0.0e+000
%NO_LO | 133 | 0.0e+000
%NO_LO | 167 | 0.0e+000
%NO_LO | 182 | 0.0e+000
%NO_LO | 223 | 0.0e+000
%NO_LO | 334 | 0.0e+000
%NO_HI | 376 | 5.6e-004
%NO_LO | 377 | 0.0e+000
%NO_LO | 378 | 0.0e+000
%NO_LO | 379 | 0.0e+000
%NO_HI | 380 | 2.1e-003
%NO_LO | 381 | 0.0e+000
%NO_LO | 628 | 0.0e+000
%NO_LO | 791 | 0.0e+000
%NO_LO | 810 | 0.0e+000
%NO_LO | 813 | 0.0e+000
%NO_LO | 820 | 0.0e+000
%NO_LO | 834 | 0.0e+000
%NO_LO | 856 | 0.0e+000
%NO_LO | 1518 | 0.0e+000
#30 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30q1t095r0TC
%TEST_NAME | Timewalk
%TEST_DATE | 30/11/2001
%TEST_TIME | 19:06:13
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 4978
%SCAN | 2
#
%HOST | ATLASPC11
%T0 | 28.0 | deg. C
%T1 | 28.0 | deg. C
#
#
# run scan charge
# 4978 2 1.25
# 4978 3 1.50
# 4978 4 1.75
# 4978 5 2.00
# 4978 6 3.00
# 4978 7 4.00
# 4978 8 5.00
# 4978 9 6.00
# 4978 10 8.00
# 4978 11 10.00
#
%00_TW | 10.9
%01_TW | 11.3
%02_TW | 11.3
%03_TW | 11.3
%04_TW | 11.2
%05_TW | 11.0
%06_TW | 10.8
%07_TW | 11.4
%08_TW | 11.6
%09_TW | 11.9
%10_TW | 11.7
%11_TW | 10.9
#
%00_TCAL | 28.7
%01_TCAL | 29.0
%02_TCAL | 30.2
%03_TCAL | 29.4
%04_TCAL | 29.5
%05_TCAL | 29.0
%06_TCAL | 29.1
%07_TCAL | 28.8
%08_TCAL | 29.3
%09_TCAL | 28.7
%10_TCAL | 28.5
%11_TCAL | 29.2
#
%DEFECTS
#name | chip | timewalk
#No defects found!

## LONGTERM
atlasPartsId | 20220170200005
eventDescription | LONGTERM-MODULE | : X=ASIC_HYBRID, MODULE
dateEntry [dd/mm/yyyy] | 07/12/2001
location [instituteCode(DB)] | KEK
personInitial | ta
problem [YES/NO] | no
pass [YES/NO] | yes
comments
dateStart [dd/mm/yyyy] | 02/12/2001
Long term test length (hours) | 24
Long term test maximum T1(xx.x) (C) | 28
Long term test minimum T1(xx.x) (C) | 25
Long term test maximum T2(xx.x) (C) | 28
Long term test minimum T2(xx.x) (C) | 27
Long term test maximum Icc (mA) | 940
Long term test minimum Icc (mA) | 930
Long term test maximum Idd (mA) | 520
Long term test minimum Idd (mA) | 470
Long term test time of first failure | no
Long term test failed ASICs | no | chip1,chip2,... (separated with comma)
RAWDATA
Time [hr] | Temp1 [C] | Temp2 [C] | Icc [mA] | Idd [mA] | every 1hr
Sat Dec 01 05:34:58 2001 | 28 | 28 | 940 | 480
Sat Dec 01 06:35:23 2001 | 25 | 27 | 940 | 480
Sat Dec 01 07:35:48 2001 | 28 | 28 | 940 | 480
Sat Dec 01 08:36:15 2001 | 28 | 28 | 940 | 480
Sat Dec 01 09:36:44 2001 | 28 | 28 | 930 | 490
Sat Dec 01 10:37:10 2001 | 28 | 27 | 930 | 470
Sat Dec 01 11:37:39 2001 | 28 | 28 | 940 | 480
Sat Dec 01 12:38:09 2001 | 28 | 27 | 930 | 470
Sat Dec 01 13:38:42 2001 | 28 | 28 | 930 | 490
Sat Dec 01 14:39:14 2001 | 28 | 27 | 940 | 470
Sat Dec 01 15:39:46 2001 | 28 | 27 | 940 | 480
Sat Dec 01 16:40:21 2001 | 28 | 27 | 940 | 480
Sat Dec 01 17:41:00 2001 | 28 | 28 | 940 | 480
Sat Dec 01 18:41:34 2001 | 28 | 28 | 940 | 480
Sat Dec 01 19:42:15 2001 | 27 | 28 | 940 | 520
Sat Dec 01 20:42:55 2001 | 28 | 27 | 930 | 520
Sat Dec 01 21:43:37 2001 | 28 | 28 | 940 | 480
Sat Dec 01 22:44:19 2001 | 28 | 27 | 930 | 480
Sat Dec 01 23:45:10 2001 | 28 | 28 | 930 | 510
Sun Dec 02 00:45:58 2001 | 28 | 28 | 940 | 470
Sun Dec 02 01:46:47 2001 | 28 | 27 | 930 | 480
Sun Dec 02 02:47:36 2001 | 28 | 28 | 940 | 480
Sun Dec 02 03:48:30 2001 | 28 | 28 | 930 | 480
Sun Dec 02 04:49:23 2001 | 28 | 28 | 940 | 470

## I_STABILITY
atlasPartsId | 20220170200005
eventDescription | I_STABILITY-LT | : X=LT
dateEntry [dd/mm/yyyy] | 07/12/2001
location [instituteCode(DB)] | KEK
personInitial | ta
problem [YES/NO] | no
pass [YES/NO] | yes
comments
biasVoltage [V] | 150
maxCurrent(20C) [microA] | 0.5852 | Current (after stabilized) corrected to 20degC
minCurrent(20C) [microA] | 0.4076 | Current (after stabilized) corrected to 20degC
deltaCurrent(20C) [microA] | 0.1776 | (max - min) corrected to 20degC
moduleRAW DATA
Time [hr] | I_total [microA] | Temp [C]
Sat Dec 01 05:34:58 2001 | 0.4838 | 28.0000 | 28.0000
Sat Dec 01 06:35:23 2001 | 0.4905 | 25.0000 | 27.0000
Sat Dec 01 07:35:48 2001 | 0.5317 | 28.0000 | 28.0000
Sat Dec 01 08:36:15 2001 | 0.4995 | 28.0000 | 28.0000
Sat Dec 01 09:36:44 2001 | 0.5009 | 28.0000 | 28.0000
Sat Dec 01 10:37:10 2001 | 0.5498 | 28.0000 | 27.0000
Sat Dec 01 11:37:39 2001 | 0.5058 | 28.0000 | 28.0000
Sat Dec 01 12:38:09 2001 | 0.4076 | 28.0000 | 27.0000
Sat Dec 01 13:38:42 2001 | 0.5123 | 28.0000 | 28.0000
Sat Dec 01 14:39:14 2001 | 0.5238 | 28.0000 | 27.0000
Sat Dec 01 15:39:46 2001 | 0.5604 | 28.0000 | 27.0000
Sat Dec 01 16:40:21 2001 | 0.5184 | 28.0000 | 27.0000
Sat Dec 01 17:41:00 2001 | 0.5217 | 28.0000 | 28.0000
Sat Dec 01 18:41:34 2001 | 0.5231 | 28.0000 | 28.0000
Sat Dec 01 19:42:15 2001 | 0.5334 | 27.0000 | 28.0000
Sat Dec 01 20:42:55 2001 | 0.5302 | 28.0000 | 27.0000
Sat Dec 01 21:43:37 2001 | 0.5387 | 28.0000 | 28.0000
Sat Dec 01 22:44:19 2001 | 0.5375 | 28.0000 | 27.0000
Sat Dec 01 23:45:10 2001 | 0.5805 | 28.0000 | 28.0000
Sun Dec 02 00:45:58 2001 | 0.5358 | 28.0000 | 28.0000
Sun Dec 02 01:46:47 2001 | 0.5449 | 28.0000 | 27.0000
Sun Dec 02 02:47:36 2001 | 0.5852 | 28.0000 | 28.0000
Sun Dec 02 03:48:30 2001 | 0.5754 | 28.0000 | 28.0000
Sun Dec 02 04:49:23 2001 | 0.5472 | 28.0000 | 28.0000

## SURVEY_XY-LT
atlasPartsId | 20220170200005
eventDescription | SURVEY_XY-LT | : X=SB, INITIAL, TC, LT, ...
date [dd/mm/yyyy] | 17/7/2001
location [instituteCode(DB)] | KEK
personInitials | TKr
problem [YES/NO] | NO
pass [YES/NO] | YES
comments | any comments
temperature(xx.x) [C] | 25.6
measurementJigID | <data>
mhx [um] | 6.7721
mhy [um] | 2.104
msx [um] | -0.73975
msy [um] | -0.32886
sepf [um] | 0.85306
sepb [um] | -2.2889
midxf [um] | 0.88542
midyf [um] | -0.64932
a1 [mrad] | 0.0075647
a2 [mrad] | 0.0015001
a3 [mrad] | -0.0045712
a4 [mrad] | -0.0089412
stereo [mrad] | -0.041724
hymxf [um] | 62.9699
hymyf [um] | 0.60188
hymaf [mrad] | 0.33191
hymxb [um] | 3.727
hymyb [um] | 5.8038
hymab [mrad] | -0.74979
conp1x [um] | 86.021
conp1y [um] | 1.8045

## SURVEY_Z-LT
atlasPartsId | 20220170200005
eventDescription | SURVEY_Z-LT | : X=INITIAL, TC, LT, ...
date [dd/mm/yyy] | 6/12/2001
location [instituteCode(DB)] | KEK
personInitial | TKr
problem [YES/NO] | NO
pass [YES/NO] | YES
comments | V2.2.0
temperature [C] | 26.5
measurementJigID | <data>
maxZlower [mm] | -0.033899
maxZupper [mm] | 0.0063817
midplaneEq | z=ax+by+c
Left a | 0.00042032
b | 0.00041927
c | 0.47291
Right a | -0.0007715
b | 0.00016382
c | 0.46822
midplaneHeight [mm] | 0.47057
moduleThickness [mm] | 1.1798
optimalMaxZerrorLower [mm] | 0.032937
optimalMaxZerrorUpper [mm] | 0.037107
optimalRmsZerrorLower [mm] | 0.0099356
optimalRmsZerrorUpper [mm] | 0.0090008
moduleConcavity x [mm] | 0.037775
y | -0.0030379
sensorSkew x [mm] | -0.023824
y | 0.018966
coolingTabThickness [mm] | 0.9349
farTabThickness [mm] | 0.95927
halfTabThickness [mm] | 0.47354
tabSkew y [mm] | -0.012184
adhesiveThicknessTotal [mm] | 0.16273
adhesiveAsymmetry [mm] | 0.0029723
loCoolingFacing a [mrad] | -0.0080295
b [mrad] | -0.27274
loCoolingFacingConcavity [mm] | 0
optimalFrameRawDATA
leftLower | x
1 | -63.1984
2 | -47.3269
3 | -31.4554
4 | -15.5838
5 | 0.28766
6 | -63.5158
7 | -47.6442
8 | -31.7727
9 | -15.9013
10 | -0.029748
11 | -63.8333
12 | -47.9617
13 | -32.0902
14 | -16.2187
15 | -0.34718
16 | -64.1508
17 | -48.2792
18 | -32.4076
19 | -16.5361
20 | -0.66459
21 | -64.4681
22 | -48.5965
23 | -32.7251
24 | -16.8536
25 | -0.98204
rightLower | x
1 | 0.98157
2 | 16.8548
3 | 32.7247
4 | 48.5962
5 | 64.4677
6 | 0.66406
7 | 16.5383
8 | 32.4073
9 | 48.2787
10 | 64.1502
11 | 0.34663
12 | 16.2199
13 | 32.0898
14 | 47.9613
15 | 63.8329
16 | 0.029318
17 | 15.9036
18 | 31.7724
19 | 47.6439
20 | 63.5154
21 | -0.28803
22 | 15.5853
23 | 31.4551
24 | 47.3266
25 | 63.198
leftUpper | x
1 | -63.1512
2 | -47.2796
3 | -31.4081
4 | -15.5364
5 | 0.33524
6 | -63.4669
7 | -47.5954
8 | -31.7236
9 | -15.852
10 | 0.019745
11 | -63.7825
12 | -47.911
13 | -32.0394
14 | -16.1676
15 | -0.29594
16 | -64.0982
17 | -48.2265
18 | -32.3548
19 | -16.4833
20 | -0.61152
21 | -64.4138
22 | -48.5421
23 | -32.6706
24 | -16.7988
25 | -0.9272
rightUpper | x
1 | 0.9275
2 | 16.7979
3 | 32.6707
4 | 48.5425
5 | 64.4143
6 | 0.61181
7 | 16.4818
8 | 32.3549
9 | 48.2268
10 | 64.0986
11 | 0.29612
12 | 16.1667
13 | 32.0394
14 | 47.9113
15 | 63.7829
16 | -0.019572
17 | 15.8505
18 | 31.7239
19 | 47.5955
20 | 63.4673
21 | -0.33505
22 | 15.5355
23 | 31.4083
24 | 47.2799
25 | 63.1518
leftLower | y
1 | -32.8277
2 | -32.508
3 | -32.1883
4 | -31.8685
5 | -31.5487
6 | -17.0563
7 | -16.7367
8 | -16.4169
9 | -16.097
10 | -15.7773
11 | -1.285
12 | -0.96508
13 | -0.64534
14 | -0.32568
15 | -0.0059087
16 | 14.4867
17 | 14.8063
18 | 15.1261
19 | 15.4457
20 | 15.7656
21 | 30.2581
22 | 30.5779
23 | 30.8976
24 | 31.2173
25 | 31.5372
rightLower | y
1 | -31.5366
2 | -31.2189
3 | -30.8971
4 | -30.5773
5 | -30.2573
6 | -15.7652
7 | -15.4485
8 | -15.1257
9 | -14.8059
10 | -14.4861
11 | 0.0063565
12 | 0.32399
13 | 0.64578
14 | 0.96563
15 | 1.2854
16 | 15.7778
17 | 16.0945
18 | 16.4173
19 | 16.737
20 | 17.0568
21 | 31.5493
22 | 31.867
23 | 32.1888
24 | 32.5085
25 | 32.8283
leftUpper | y
1 | -32.8215
2 | -32.5038
3 | -32.1859
4 | -31.8684
5 | -31.5506
6 | -17.0507
7 | -16.7328
8 | -16.4151
9 | -16.0975
10 | -15.7797
11 | -1.2797
12 | -0.96198
13 | -0.64424
14 | -0.32653
15 | -0.0089433
16 | 14.4911
17 | 14.8088
18 | 15.1266
19 | 15.4443
20 | 15.7618
21 | 30.262
22 | 30.5796
23 | 30.8974
24 | 31.2151
25 | 31.5327
rightUpper | y
1 | -31.5331
2 | -31.2134
3 | -30.8978
4 | -30.5802
5 | -30.2627
6 | -15.7623
7 | -15.4415
8 | -15.1269
9 | -14.8094
10 | -14.4918
11 | 0.0086614
12 | 0.32828
13 | 0.6439
14 | 0.96146
15 | 1.279
16 | 15.7795
17 | 16.1001
18 | 16.4146
19 | 16.7323
20 | 17.0499
21 | 31.5503
22 | 31.8699
23 | 32.1855
24 | 32.5031
25 | 32.8207
leftLower | z
1 | 0.085672
2 | 0.040264
3 | 0.02132
4 | 0.017413
5 | 0.014641
6 | 0.059753
7 | 0.016835
8 | 0.0046465
9 | 0.0017775
10 | 0.0011064
11 | 0.053225
12 | 0.011172
13 | 0.0027896
14 | -0.0013105
15 | 0.0028465
16 | 0.062274
17 | 0.016596
18 | 0.00028064
19 | -0.0055184
20 | -0.0092724
21 | 0.088034
22 | 0.038398
23 | 0.013426
24 | 0.0053227
25 | 0.0051447
rightLower | z
1 | 0.010082
2 | 0.013047
3 | 0.016011
4 | 0.040491
5 | 0.089078
6 | -0.00029857
7 | -0.0017122
8 | -0.0031272
9 | 0.014583
10 | 0.05847
11 | 0.003267
12 | -0.0015999
13 | -0.0064677
14 | 0.00529
15 | 0.045976
16 | -0.00040236
17 | -0.0027189
18 | -0.0050371
19 | 0.0088914
20 | 0.051352
21 | 0.016506
22 | 0.014209
23 | 0.011912
24 | 0.032804
25 | 0.078282
leftUpper | z
1 | -0.096906
2 | -0.046354
3 | -0.018267
4 | -0.0083698
5 | -0.0086222
6 | -0.066504
7 | -0.018117
8 | -0.001727
9 | 0.0031885
10 | -0.001048
11 | -0.054544
12 | -0.010508
13 | 0.0018363
14 | 0.0013508
15 | -0.0055448
16 | -0.05871
17 | -0.012201
18 | 0.0019525
19 | 0.001403
20 | -0.0023146
21 | -0.082872
22 | -0.034304
23 | -0.011745
24 | -0.0072928
25 | -0.010609
rightUpper | z
1 | -0.012781
2 | -0.013714
3 | -0.014646
4 | -0.030191
5 | -0.070213
6 | -0.0013981
7 | -0.0032093
8 | -0.0050207
9 | -0.015161
10 | -0.052826
11 | -0.0068824
12 | -0.0048226
13 | -0.0027629
14 | -0.012378
15 | -0.050219
16 | 0.0045253
17 | 0.0012321
18 | -0.0020611
19 | -0.01669
20 | -0.057433
21 | 0.0007701
22 | -0.0037311
23 | -0.0082324
24 | -0.033256
25 | -0.077779

## ELECTRICAL-LT
atlasPartsId | 20220170200005
eventDescription | ELECTRICAL-LT |  | : X=INITIAL, TC, LT, ANNEAL, ...
date [dd/mm/yyyy] | 05/12/2001
location [instituteCode(DB)] | KEK
personInitial | ta
problem [YES/NO] | no
pass [YES/NO] | yes
comments
#
# | http://hepwww.rl.ac.uk/atlas-sct/documents/dev/Electrical_Tests_301.doc
# | N.B. Some revisions have been incorporated here, but not yet in the reference,
# | namely that CONFIGURATION will be stored as an ASCII file and
# | and that the RUN field will be used to store both RUN and SCAN number.
#
#
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mLT
%TEST_NAME | FullBypassTest
%TEST_DATE | 37023
%TEST_TIME | 0.65697
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 5070
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 27 | deg. C
%T1 | 27 | deg. C
#
#NConfigs 37
#Vdd = 4.0 - 3.5
#
%M0->S1: | 3.5 V | (Minimal tested)
%M0->S2: | 3.5 V | (Minimal tested)
%S1->S2: | 3.5 V | (Minimal tested)
%S1->S3: | 3.5 V | (Minimal tested)
%S2->S3: | 3.5 V | (Minimal tested)
%S2->S4: | 3.5 V | (Minimal tested)
%S3->S4: | 3.5 V | (Minimal tested)
%S3->E5: | 3.5 V | (Minimal tested)
%S4->E5: | 3.5 V | (Minimal tested)
%E5->S9: | 3.5 V | (Minimal tested)
%M8->S9: | 3.5 V | (Minimal tested)
%M8->S10: | 3.5 V | (Minimal tested)
%S9->S10: | 3.5 V | (Minimal tested)
%S9->S11: | 3.5 V | (Minimal tested)
%S10->S11: | 3.5 V | (Minimal tested)
%S10->S12: | 3.5 V | (Minimal tested)
%S11->S12: | 3.5 V | (Minimal tested)
%S11->E13: | 3.5 V | (Minimal tested)
%S12->E13: | 3.5 V | (Minimal tested)
%E13->S1: | 3.5 V | (Minimal tested)
#
%DEFECTS
#name | bypass_link | Vdd_min
#No defects found!
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mLT
%TEST_NAME | RedundancyTest
%TEST_DATE | 37023
%TEST_TIME | 0.65715
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | NO
%PROBLEM | NO
#
%RUN | 5070
%SCAN | 7
#
%HOST | ATLASPC11
%T0 | 27 | deg. C
%T1 | 26 | deg. C
#
#RedundancyTest Summary - not for the database
#chip pass good stuck dead adr0 adr1 com0 com1 com
# 0 0 0 0 0 0 128 0 0 0
# 1 0 0 0 0 0 128 0 0 0
# 2 0 0 0 0 0 128 0 0 0
# 3 0 0 0 0 0 128 0 0 0
# 4 0 0 0 0 0 128 0 0 0
# 5 0 0 0 0 0 128 0 0 0
# 6 0 0 0 0 0 128 0 0 0
# 7 0 0 0 0 0 128 0 0 0
# 8 0 0 0 0 0 128 0 0 0
# 9 0 0 0 0 0 128 0 0 0
#10 0 0 0 0 0 128 0 0 0
#11 0 0 0 0 0 128 0 0 0
#
%DEFECTS
#name | chip
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30
%TEST_NAME | StrobeDelay
%TEST_DATE | 03/12/2001
%TEST_TIME | 13:44:19
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 5031
%SCAN | 2
#
%HOST | ATLASPC11
%T0 | 25.0 | deg. C
%T1 | 27.0 | deg. C
#
#
#parameter | value
%00_DELAY | 11
%01_DELAY | 12
%02_DELAY | 12
%03_DELAY | 11
%04_DELAY | 12
%05_DELAY | 12
%06_DELAY | 12
%07_DELAY | 11
%08_DELAY | 12
%09_DELAY | 12
%10_DELAY | 12
%11_DELAY | 14
#
%DEFECTS
#name | chip
#No defects found!
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30
%TEST_NAME | ResponseCurve
%TEST_DATE | 03/12/2001
%TEST_TIME | 13:54:05
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 5032
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 27.0 | deg. C
%T1 | 27.0 | deg. C
#
#
%N_SCANS | 3
%Q_00 | 1.50 fC
%Q_01 | 2.00 fC
%Q_02 | 2.50 fC
%RC_FIT | 4 | (linear fit)
#
%ModuleResponseCurvePlot
# |  | func | p0 | p1 | p2
Chip 0: | 4 | 20.97 | 55.77 | 0.00
Chip 1: | 4 | 19.14 | 55.69 | 0.00
Chip 2: | 4 | 15.07 | 54.42 | 0.00
Chip 3: | 4 | 19.32 | 55.76 | 0.00
Chip 4: | 4 | 18.78 | 57.55 | 0.00
Chip 5: | 4 | 22.74 | 54.34 | 0.00
Chip 6: | 4 | 20.16 | 55.37 | 0.00
Chip 7: | 4 | 16.84 | 56.81 | 0.00
Chip 8: | 4 | 17.74 | 55.18 | 0.00
Chip 9: | 4 | 16.81 | 54.80 | 0.00
Chip 10: | 4 | 15.06 | 55.44 | 0.00
Chip 11: | 4 | 19.96 | 54.10 | 0.00
#
%ModuleGainOffsetNoise
# at 2.00fC
# |  | vt50 | rms | gain | rms | offset | rms | outnse | innse | rms | comment
Chip 0: | 132.1 | 7.63 | 55.8 | 0.99 | 21.4 | 7.19 | 12.76 | 1431 | 29 | OK
Chip 1: | 130.5 | 7.65 | 55.7 | 1.04 | 19.3 | 7.75 | 12.72 | 1428 | 28 | OK
Chip 2: | 125.3 | 6.99 | 54.3 | 3.70 | 15.3 | 6.28 | 13.12 | 1424 | 107 | OK
Chip 3: | 130.9 | 7.59 | 55.8 | 1.10 | 19.5 | 6.98 | 13.06 | 1464 | 29 | OK
Chip 4: | 134.1 | 8.11 | 57.5 | 1.16 | 19.1 | 7.09 | 13.23 | 1437 | 25 | OK
Chip 5: | 131.3 | 8.44 | 54.3 | 1.00 | 23.1 | 7.98 | 12.86 | 1479 | 35 | OK
Chip 6: | 130.9 | 7.62 | 55.4 | 1.05 | 20.2 | 7.64 | 12.56 | 1418 | 33 | OK
Chip 7: | 130.6 | 6.58 | 56.8 | 0.99 | 16.8 | 6.66 | 13.18 | 1450 | 27 | OK
Chip 8: | 128.0 | 7.89 | 55.2 | 1.03 | 17.9 | 7.26 | 12.83 | 1453 | 31 | OK
Chip 9: | 126.7 | 8.44 | 54.8 | 1.01 | 17.6 | 7.76 | 12.87 | 1469 | 31 | OK
Chip 10: | 126.0 | 8.15 | 55.4 | 1.16 | 15.3 | 7.65 | 12.77 | 1440 | 31 | OK
Chip 11: | 128.0 | 7.85 | 54.1 | 1.13 | 20.0 | 7.42 | 13.23 | 1529 | 39 | OK
#
#BadChannelSummary - not for the database
# at 2.00fC
# |  | lost | dodgy | dead | stuck | ineff | unbon | lo_gn | hi_gn | lo_off | hi_off | partbon | hi_nse
#Chip 0: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 1: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 2: | 2 | 2 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 2
#Chip 3: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 4: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 5: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 6: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 7: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 8: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 9: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 10: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 11: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Link 0: | 2 | 2 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 2
#Link 1: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Link 2: | 2 | 2 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 2
#
%DEFECTS
#name | channel | value
%NOISY | 377 | 12801.71
%NOISY | 378 | 12195.17
%UNBONDED | 379 | 592.58
%UNBONDED | 381 | 628.32
#4 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200003mb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 03/12/2001
%TEST_TIME | 14:28:16
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | NO
%PROBLEM | NO
#
%RUN | 5033
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 29.0 | deg. C
%T1 | 28.0 | deg. C
#
%TRIM_TYPE | 0
%TRIM_Q | 10.0
%N_SCANS | 16
%T_00 | 0
%T_01 | 1
%T_02 | 2
%T_03 | 3
%T_04 | 4
%T_05 | 5
%T_06 | 6
%T_07 | 7
%T_08 | 8
%T_09 | 9
%T_10 | 10
%T_11 | 11
%T_12 | 12
%T_13 | 13
%T_14 | 14
%T_15 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005mb220s30_tr0_20011203.trim
%MASK_FILE | D:\sctvar\results\0170200005mb220s30_tr0_20011203.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 0 | 95.0 | 128 | 95.1 | 1.33 | -19.3 | 2.0 | 3.87 | 0.17
Chip 1: | 0 | 95.0 | 128 | 95.1 | 1.22 | -20.5 | 2.5 | 3.56 | 0.18
Chip 2: | 0 | 95.0 | 127 | 95.0 | 1.10 | -22.9 | 3.1 | 3.05 | 0.26
Chip 3: | 0 | 95.0 | 128 | 95.0 | 1.34 | -19.2 | 2.0 | 3.83 | 0.16
Chip 4: | 0 | 95.0 | 127 | 95.0 | 1.12 | -21.2 | 2.2 | 3.56 | 0.19
Chip 5: | 0 | 95.0 | 128 | 95.2 | 1.08 | -21.9 | 2.7 | 3.39 | 0.12
Chip 6: | 0 | 95.0 | 128 | 95.3 | 1.17 | -20.5 | 2.3 | 3.59 | 0.16
Chip 7: | 0 | 95.0 | 128 | 94.8 | 1.19 | -20.4 | 2.2 | 3.57 | 0.18
Chip 8: | 0 | 95.0 | 128 | 95.1 | 1.17 | -20.5 | 2.4 | 3.48 | 0.15
Chip 9: | 0 | 95.0 | 128 | 95.0 | 1.16 | -20.9 | 2.6 | 3.38 | 0.15
Chip 10: | 0 | 95.0 | 128 | 95.0 | 1.21 | -19.8 | 2.4 | 3.49 | 0.14
Chip 11: | 0 | 95.0 | 127 | 95.1 | 1.59 | -17.5 | 1.9 | 4.14 | 0.36
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -11.90
%TR_STEP | 127 | 4.42
%TR_STEP | 165 | 4.75
%TR_OFFSET | 209 | -28.49
%TR_OFFSET | 378 | -13.37
%TR_OFFSET | 452 | -11.93
%TR_STEP | 511 | 4.44
%TR_STEP | 639 | 4.25
%TR_OFFSET | 171 | -27.15
%TR_OFFSET | 210 | -27.36
%TR_STEP | 251 | 4.19
%TR_STEP | 255 | 4.19
%TR_OFFSET | 585 | -27.95
%TR_OFFSET | 750 | -25.21
%TR_STEP | 765 | 5.27
#15 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200003mb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 03/12/2001
%TEST_TIME | 14:28:16
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 5033
%SCAN | 17
#
%HOST | ATLASPC11
%T0 | 29.0 | deg. C
%T1 | 28.0 | deg. C
#
%TRIM_TYPE | 1
%TRIM_Q | 10.0
%N_SCANS | 4
%T_00 | 3
%T_01 | 7
%T_02 | 11
%T_03 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005mb220s30_tr1_20011203.trim
%MASK_FILE | D:\sctvar\results\0170200005mb220s30_tr1_20011203.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 1 | 120.0 | 128 | 120.2 | 2.12 | -10.4 | 1.0 | 7.20 | 0.30
Chip 1: | 1 | 120.0 | 128 | 120.1 | 1.90 | -10.9 | 1.5 | 6.74 | 0.43
Chip 2: | 1 | 120.0 | 128 | 120.1 | 1.91 | -11.2 | 1.4 | 6.18 | 0.42
Chip 3: | 1 | 120.0 | 128 | 120.0 | 2.05 | -10.3 | 1.1 | 7.21 | 0.27
Chip 4: | 1 | 120.0 | 128 | 119.6 | 2.00 | -11.1 | 1.2 | 6.84 | 0.32
Chip 5: | 1 | 120.0 | 128 | 120.1 | 1.87 | -11.6 | 1.4 | 6.48 | 0.23
Chip 6: | 1 | 120.0 | 128 | 119.8 | 1.88 | -10.8 | 1.2 | 6.85 | 0.28
Chip 7: | 1 | 120.0 | 128 | 120.1 | 2.01 | -10.4 | 1.2 | 6.99 | 0.31
Chip 8: | 1 | 120.0 | 128 | 120.1 | 1.97 | -10.6 | 1.3 | 6.73 | 0.28
Chip 9: | 1 | 120.0 | 128 | 120.0 | 1.86 | -10.9 | 1.4 | 6.52 | 0.26
Chip 10: | 1 | 120.0 | 128 | 120.1 | 1.76 | -10.2 | 1.3 | 6.78 | 0.26
Chip 11: | 1 | 120.0 | 128 | 120.5 | 2.41 | -9.4 | 1.0 | 7.74 | 0.58
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -6.58
%TR_STEP | 165 | 10.54
%TR_OFFSET | 378 | -6.42
%TR_OFFSET | 452 | -6.34
%TR_STEP | 511 | 8.07
%TR_STEP | 639 | 7.85
%TR_OFFSET | 691 | -7.36
%TR_OFFSET | 734 | -7.15
%TR_OFFSET | 171 | -14.03
%TR_OFFSET | 210 | -14.11
%TR_STEP | 251 | 8.00
%TR_OFFSET | 585 | -14.04
%TR_OFFSET | 750 | -13.90
%TR_STEP | 765 | 9.60
#14 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200003mb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 03/12/2001
%TEST_TIME | 14:28:16
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 5033
%SCAN | 21
#
%HOST | ATLASPC11
%T0 | 29.0 | deg. C
%T1 | 28.0 | deg. C
#
%TRIM_TYPE | 2
%TRIM_Q | 10.0
%N_SCANS | 4
%T_00 | 3
%T_01 | 7
%T_02 | 11
%T_03 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005mb220s30_tr2_20011203.trim
%MASK_FILE | D:\sctvar\results\0170200005mb220s30_tr2_20011203.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 2 | 140.0 | 128 | 139.9 | 3.22 | -7.0 | 0.6 | 10.70 | 0.46
Chip 1: | 2 | 140.0 | 128 | 140.5 | 2.61 | -7.3 | 1.0 | 10.03 | 0.63
Chip 2: | 2 | 140.0 | 128 | 139.5 | 2.78 | -7.4 | 0.9 | 9.32 | 0.59
Chip 3: | 2 | 140.0 | 128 | 140.0 | 3.19 | -6.9 | 0.7 | 10.76 | 0.41
Chip 4: | 2 | 140.0 | 128 | 140.1 | 3.05 | -7.4 | 0.8 | 10.16 | 0.44
Chip 5: | 2 | 140.0 | 128 | 139.8 | 2.73 | -7.7 | 0.9 | 9.69 | 0.34
Chip 6: | 2 | 140.0 | 128 | 139.7 | 2.73 | -7.3 | 0.7 | 10.19 | 0.42
Chip 7: | 2 | 140.0 | 128 | 139.9 | 3.18 | -6.9 | 0.8 | 10.51 | 0.43
Chip 8: | 2 | 140.0 | 128 | 140.1 | 2.82 | -7.1 | 0.8 | 9.98 | 0.40
Chip 9: | 2 | 140.0 | 128 | 140.3 | 2.85 | -7.3 | 0.9 | 9.72 | 0.39
Chip 10: | 2 | 140.0 | 128 | 139.7 | 2.95 | -6.9 | 0.8 | 10.04 | 0.39
Chip 11: | 2 | 140.0 | 128 | 139.7 | 3.26 | -6.5 | 0.7 | 11.20 | 0.77
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -4.53
%TR_STEP | 165 | 15.81
%TR_OFFSET | 378 | -4.44
%TR_OFFSET | 452 | -4.10
%TR_STEP | 511 | 12.02
%TR_STEP | 639 | 11.62
%TR_OFFSET | 691 | -4.84
%TR_OFFSET | 734 | -5.05
%TR_OFFSET | 171 | -9.27
%TR_OFFSET | 210 | -9.40
%TR_STEP | 251 | 11.85
%TR_OFFSET | 585 | -9.42
%TR_STEP | 627 | 11.22
%TR_OFFSET | 750 | -9.59
%TR_STEP | 765 | 13.69
%TR_STEP | 766 | 13.57
#16 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200003mb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 03/12/2001
%TEST_TIME | 14:28:16
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 5033
%SCAN | 25
#
%HOST | ATLASPC11
%T0 | 29.0 | deg. C
%T1 | 28.0 | deg. C
#
%TRIM_TYPE | 3
%TRIM_Q | 10.0
%N_SCANS | 4
%T_00 | 3
%T_01 | 7
%T_02 | 11
%T_03 | 15
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005mb220s30_tr3_20011203.trim
%MASK_FILE | D:\sctvar\results\0170200005mb220s30_tr3_20011203.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 3 | 160.0 | 128 | 159.3 | 4.22 | -5.3 | 0.5 | 14.11 | 0.56
Chip 1: | 3 | 160.0 | 128 | 160.2 | 3.55 | -5.4 | 0.8 | 13.36 | 0.81
Chip 2: | 3 | 160.0 | 128 | 160.0 | 3.65 | -5.6 | 0.7 | 12.43 | 0.72
Chip 3: | 3 | 160.0 | 128 | 160.1 | 4.02 | -5.2 | 0.5 | 14.18 | 0.49
Chip 4: | 3 | 160.0 | 128 | 160.5 | 3.81 | -5.7 | 0.6 | 13.39 | 0.58
Chip 5: | 3 | 160.0 | 128 | 160.1 | 3.67 | -5.7 | 0.6 | 12.99 | 0.45
Chip 6: | 3 | 160.0 | 128 | 159.8 | 3.73 | -5.4 | 0.6 | 13.55 | 0.53
Chip 7: | 3 | 160.0 | 128 | 160.0 | 4.00 | -5.3 | 0.6 | 13.85 | 0.56
Chip 8: | 3 | 160.0 | 128 | 159.9 | 3.78 | -5.4 | 0.6 | 13.24 | 0.52
Chip 9: | 3 | 160.0 | 128 | 160.1 | 3.70 | -5.4 | 0.7 | 12.91 | 0.50
Chip 10: | 3 | 160.0 | 128 | 160.2 | 3.99 | -5.2 | 0.6 | 13.31 | 0.50
Chip 11: | 3 | 160.0 | 128 | 159.5 | 4.11 | -4.9 | 0.5 | 14.79 | 0.92
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -3.29
%TR_STEP | 165 | 20.87
%TR_STEP | 258 | 14.62
%TR_OFFSET | 378 | -3.07
%TR_OFFSET | 452 | -3.27
%TR_STEP | 511 | 15.75
%TR_OFFSET | 691 | -3.77
%TR_OFFSET | 734 | -3.76
%TR_OFFSET | 171 | -7.09
%TR_OFFSET | 210 | -7.24
%TR_STEP | 251 | 15.54
%TR_OFFSET | 585 | -7.04
%TR_STEP | 623 | 14.91
%TR_OFFSET | 750 | -7.13
%TR_STEP | 765 | 17.77
#15 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200003mb220s30
%TEST_NAME | TrimScan
%TEST_DATE | 03/12/2001
%TEST_TIME | 14:28:16
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
%PASS | YES
%PROBLEM | NO
#
%RUN | 5033
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 29.0 | deg. C
%T1 | 28.0 | deg. C
#
%TRIM_TYPE | -1
%TRIM_Q | 10.0
%TRIM_ALGORITHM | 0
#(trim whole module)
#
%TRIM_FILE | D:\sctvar\results\0170200005mb220s30_tr-1_20011203.trim
%MASK_FILE | D:\sctvar\results\0170200005mb220s30_tr-1_20011203.mask
#
%TrimSummary
#
# |  | range | target | ntrim | vt50 | vt50rms | tr_off | off_rms | tr_step | step_rms
Chip 0: | 0 | 102.5 | 128 | 102.1 | 1.06 | -19.3 | 2.0 | 3.87 | 0.17
Chip 1: | 0 | 102.5 | 128 | 102.0 | 1.02 | -20.5 | 2.5 | 3.56 | 0.18
Chip 2: | 1 | 102.5 | 128 | 102.0 | 1.90 | -11.2 | 1.4 | 6.18 | 0.42
Chip 3: | 1 | 102.5 | 128 | 101.9 | 2.04 | -10.3 | 1.1 | 7.21 | 0.27
Chip 4: | 0 | 102.5 | 128 | 102.0 | 1.02 | -21.2 | 2.2 | 3.56 | 0.19
Chip 5: | 0 | 102.5 | 128 | 102.0 | 0.95 | -21.9 | 2.7 | 3.39 | 0.12
Chip 6: | 0 | 102.5 | 128 | 102.1 | 1.05 | -20.5 | 2.3 | 3.59 | 0.16
Chip 7: | 0 | 102.5 | 128 | 102.0 | 1.06 | -20.4 | 2.2 | 3.57 | 0.18
Chip 8: | 0 | 102.5 | 128 | 102.0 | 0.96 | -20.5 | 2.4 | 3.48 | 0.15
Chip 9: | 1 | 102.5 | 128 | 101.9 | 1.83 | -10.9 | 1.4 | 6.52 | 0.26
Chip 10: | 0 | 102.5 | 128 | 102.0 | 1.01 | -19.8 | 2.4 | 3.49 | 0.14
Chip 11: | 1 | 102.5 | 128 | 101.6 | 2.29 | -9.4 | 1.0 | 7.74 | 0.58
#
%DEFECTS
#name | channel | value
%TR_OFFSET | 59 | -11.90
%TR_STEP | 127 | 4.42
%TR_STEP | 165 | 4.75
%TR_OFFSET | 209 | -28.49
%TR_OFFSET | 378 | -6.42
%TR_OFFSET | 452 | -6.34
%TR_STEP | 511 | 8.07
%TR_STEP | 639 | 4.25
%TR_OFFSET | 171 | -27.15
%TR_OFFSET | 210 | -27.36
%TR_STEP | 251 | 4.19
%TR_STEP | 255 | 4.19
%TR_OFFSET | 585 | -27.95
%TR_OFFSET | 750 | -13.90
%TR_STEP | 765 | 9.60
#15 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30q1t095r0LT
%TEST_NAME | ResponseCurve
%TEST_DATE | 03/12/2001
%TEST_TIME | 15:20:07
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 5035
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 28.0 | deg. C
%T1 | 28.0 | deg. C
#
#
%N_SCANS | 10
%Q_00 | 0.50 fC
%Q_01 | 0.75 fC
%Q_02 | 1.00 fC
%Q_03 | 1.25 fC
%Q_04 | 1.50 fC
%Q_05 | 2.00 fC
%Q_06 | 3.00 fC
%Q_07 | 4.00 fC
%Q_08 | 6.00 fC
%Q_09 | 8.00 fC
%RC_FIT | 3 | (exponential fit)
#
%ModuleResponseCurvePlot
# |  | func | p0 | p1 | p2
Chip 0: | 3 | 1645.42 | 7.10 | -785.46
Chip 1: | 3 | 1601.75 | 7.04 | -762.23
Chip 2: | 3 | 1413.56 | 6.43 | -666.44
Chip 3: | 3 | 1609.36 | 7.09 | -765.84
Chip 4: | 3 | 1586.90 | 6.84 | -755.49
Chip 5: | 3 | 1508.28 | 6.69 | -714.92
Chip 6: | 3 | 1551.28 | 6.84 | -736.57
Chip 7: | 3 | 1626.30 | 7.16 | -774.22
Chip 8: | 3 | 1528.00 | 6.81 | -724.49
Chip 9: | 3 | 1540.92 | 6.97 | -730.14
Chip 10: | 3 | 1495.52 | 6.77 | -707.04
Chip 11: | 3 | 1810.27 | 8.15 | -865.16
#
%ModuleGainOffsetNoise
# at 2.00fC
# |  | vt50 | rms | gain | rms | offset | rms | outnse | innse | rms | comment
Chip 0: | 151.6 | 1.59 | 56.7 | 1.04 | 37.4 | 1.66 | 12.45 | 1374 | 27 | OK
Chip 1: | 151.4 | 1.49 | 55.7 | 1.05 | 38.7 | 1.72 | 12.36 | 1386 | 27 | OK
Chip 2: | 150.2 | 1.45 | 53.5 | 1.10 | 40.7 | 1.55 | 12.21 | 1425 | 109 | OK
Chip 3: | 151.5 | 1.71 | 55.5 | 1.17 | 39.0 | 1.74 | 12.64 | 1423 | 29 | OK
Chip 4: | 153.4 | 1.74 | 56.6 | 1.15 | 38.2 | 1.69 | 12.91 | 1425 | 22 | OK
Chip 5: | 150.5 | 1.49 | 54.8 | 1.10 | 39.6 | 1.59 | 12.34 | 1407 | 32 | OK
Chip 6: | 151.4 | 1.56 | 55.3 | 1.06 | 39.4 | 1.47 | 12.06 | 1363 | 29 | OK
Chip 7: | 152.2 | 1.48 | 55.6 | 0.94 | 39.2 | 1.52 | 12.86 | 1446 | 32 | OK
Chip 8: | 150.8 | 1.46 | 54.8 | 1.03 | 39.8 | 1.46 | 12.52 | 1429 | 29 | OK
Chip 9: | 150.3 | 1.36 | 54.1 | 0.95 | 40.4 | 1.55 | 12.35 | 1428 | 28 | OK
Chip 10: | 150.9 | 1.38 | 54.0 | 1.09 | 40.8 | 1.68 | 12.44 | 1439 | 29 | OK
Chip 11: | 150.0 | 1.96 | 54.6 | 1.25 | 40.2 | 1.81 | 12.77 | 1463 | 36 | OK
#
#BadChannelSummary - not for the database
# at 2.00fC
# |  | lost | dodgy | dead | stuck | ineff | unbon | lo_gn | hi_gn | lo_off | hi_off | partbon | hi_nse
#Chip 0: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 1: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 2: | 4 | 0 | 2 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 3: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 4: | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 5: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 6: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 7: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 8: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 9: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 10: | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Chip 11: | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Link 0: | 5 | 0 | 3 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0
#Link 1: | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
#Link 2: | 6 | 0 | 4 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0
#
%DEFECTS
#name | channel | value
%DEAD | 377
%DEAD | 378
%UNBONDED | 379 | 578.08
%UNBONDED | 381 | 629.68
%DEAD | 628
%DEAD | 1518
#6 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30q1t095r0LT
%TEST_NAME | NoiseOccupancy
%TEST_DATE | 03/12/2001
%TEST_TIME | 15:40:41
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 5036
%SCAN | 1
#
%HOST | ATLASPC11
%T0 | 28.0 | deg. C
%T1 | 27.0 | deg. C
#
#
%IccMax: 940
%IddMax: 760
%IccTyp: 940
%IddTyp: 490
#
%NOSummary
#
# |  | Offset | MeanOcc | RMSOcc | EstENC
Chip 0: | 21.8 | 3.3e-006 | 2.6e-006 | 1687
Chip 1: | 23.1 | 5.3e-006 | 4.4e-006 | 1716
Chip 2: | -0.9 | 2.7e-005 | 1.6e-004 | 2017
Chip 3: | 18.9 | 2.3e-005 | 2.1e-005 | 1870
Chip 4: | 14.7 | 3.5e-005 | 2.5e-005 | 1909
Chip 5: | 14.1 | 5.6e-005 | 6.3e-005 | 2050
Chip 6: | 21.9 | 7.8e-006 | 1.0e-005 | 1750
Chip 7: | 20.7 | 1.3e-005 | 7.7e-006 | 1770
Chip 8: | 17.9 | 1.6e-005 | 1.1e-005 | 1841
Chip 9: | 19.5 | 2.1e-005 | 1.2e-005 | 1856
Chip 10: | 18.8 | 2.4e-005 | 1.4e-005 | 1882
Chip 11: | 8.8 | 8.7e-005 | 7.6e-005 | 2152
#
%DEFECTS
#name | chan
%NO_LO | 8 | 0.0e+000
%NO_LO | 26 | 0.0e+000
%NO_LO | 31 | 0.0e+000
%NO_LO | 38 | 0.0e+000
%NO_LO | 52 | 0.0e+000
%NO_LO | 56 | 0.0e+000
%NO_LO | 57 | 0.0e+000
%NO_LO | 60 | 0.0e+000
%NO_LO | 61 | 0.0e+000
%NO_LO | 66 | 0.0e+000
%NO_LO | 70 | 0.0e+000
%NO_LO | 73 | 0.0e+000
%NO_LO | 80 | 0.0e+000
%NO_LO | 88 | 0.0e+000
%NO_LO | 100 | 0.0e+000
%NO_LO | 122 | 0.0e+000
%NO_LO | 124 | 0.0e+000
%NO_LO | 143 | 0.0e+000
%NO_LO | 147 | 0.0e+000
%NO_LO | 167 | 0.0e+000
%NO_LO | 182 | 0.0e+000
%NO_LO | 195 | 0.0e+000
%NO_LO | 202 | 0.0e+000
%NO_LO | 204 | 0.0e+000
%NO_LO | 207 | 0.0e+000
%NO_LO | 218 | 0.0e+000
%NO_LO | 220 | 0.0e+000
%NO_LO | 342 | 0.0e+000
%NO_LO | 377 | 0.0e+000
%NO_LO | 378 | 0.0e+000
%NO_LO | 379 | 0.0e+000
%NO_HI | 380 | 1.8e-003
%NO_LO | 381 | 0.0e+000
%NO_LO | 628 | 0.0e+000
%NO_LO | 840 | 0.0e+000
%NO_LO | 857 | 0.0e+000
%NO_LO | 868 | 0.0e+000
%NO_LO | 1518 | 0.0e+000
#38 defects found
#
%NEW_TEST
#
%SERIAL_NO | 0170200005mb220s30q1t095r0LT
%TEST_NAME | Timewalk
%TEST_DATE | 03/12/2001
%TEST_TIME | 16:10:51
%LOCN_NAME | UNKNOWN
%INITLS | Administrator
#
%PASS | YES
%PROBLEM | NO
#
%RUN | 5037
%SCAN | 2
#
%HOST | ATLASPC11
%T0 | 28.0 | deg. C
%T1 | 28.0 | deg. C
#
#
# run scan charge
# 5037 2 1.25
# 5037 3 1.50
# 5037 4 1.75
# 5037 5 2.00
# 5037 6 3.00
# 5037 7 4.00
# 5037 8 5.00
# 5037 9 6.00
# 5037 10 8.00
# 5037 11 10.00
#
%00_TW | 10.7
%01_TW | 11.1
%02_TW | 11.2
%03_TW | 11.1
%04_TW | 11.0
%05_TW | 10.9
%06_TW | 10.6
%07_TW | 11.1
%08_TW | 11.4
%09_TW | 11.7
%10_TW | 11.5
%11_TW | 10.7
#
%00_TCAL | 28.8
%01_TCAL | 29.1
%02_TCAL | 30.3
%03_TCAL | 29.4
%04_TCAL | 29.6
%05_TCAL | 29.1
%06_TCAL | 29.2
%07_TCAL | 28.9
%08_TCAL | 29.4
%09_TCAL | 28.8
%10_TCAL | 28.5
%11_TCAL | 29.2
#
%DEFECTS
#name | chip | timewalk
#No defects found!

## SUMMARY
atlasPartsId | 20220170200005
eventDescription | MODULE-SUMMARY
dateEntry [dd/mm/yyyy] | 07/12/2001
location [instituteCode(DB)] | KEK
personInitial | ta
problem [YES/NO] | no
pass [YES/NO] | yes
comments
last I_V I_150V [microA] | 0.1428
I_350V [microA] | 0.1959
I_500V [microA] | 0.2253
lastLT deltaI [microA] | 0.1776
maxZupper [mm] | 0.0063817
maxZlower [mm] | -0.033899
nDefectiveChannels | 6
defectiveChannels | 377, 378, 379, 381, 628, 1518 |  | ch1,ch2,ch3,..., in ONE cell, separated with comma
averageNoiseOccupancy (1fC) | 2.6533e-05 | : of good channels
lastMetrologyDate [dd/mm/yyyy] | 06/12/2001
lastI_Vdate[dd/mm/yyyy] | 06/12/2001